## Sheet1
Task No | Task | Bucket Name | Sub-bucket | Status | Progress | Assignees | Internal unit Testing | Priority | Completed Date | Created Date | Start date | End date | Completed By | Completed Checklist Items | Checklist Items | Remarks | Attachment (Proof) | Item Type | Path
1 | 50 Hz Notch filter Feasibility and implementation if possible | Genx12 sustenance |  | Completed | 100 | Anubhav Chaudhary;sanjana sampige;somesh | false | Low |  | 2025-03-13 01:00 |  | 2025-05-26 |  | 4/6 | 5. Update the UI part with keyboard and touch functionality.;2. Will test and compare the filter working with counterpart machine.;3. Integrate 60 Hz filter with main code.;6. Update Filter Code as it did not show good response Impulse Response ;4. Test the CPU cycles consumed by filters(how much time it takes).;1.Integrated 50 hz filter generated from APP before 17/03/2025. |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
2 | GENX3 UART DMA Check | Genx3 Sustenance |  | Completed | 100 | Anubhav Chaudhary;Someshwaran S | N.A. | Medium |  | 2025-03-07 | 2025-02-07 | 2025-05-27 |  |  |  | Checking with 256k baud rate -Live transmission getting stopped;tried to transfer all the stored records (495) 1st time 438 records transferred, 2nd time all the 495 records tarnsferred 256K baud rate is working in Record transfer ,but in live transfer some issue observed. tried with 921600 baud rate working fine in record transfer. Found some issues in live transfer |  | Item | sites/ECGTeam761/Lists/Genx sustenance
3 | 4_11 NiBP Limits: The limits for NIBP should be configurable and not to be fixed or auto corrected to upper limits. | Genx12 sustenance |  | Completed | 100 | Mahesh DE | false | Medium |  | 2025-04-09 01:00 | [5 empty] | Ref: Re: Observations from ECG Camp – GenX12i+ Performance and Improvement Suggestions 7/03/2025 ¦ Solution: ¦ ¦ --- ¦ ¦ We can increase it to 999/999 if required. ¦ |  | Item | sites/ECGTeam761/Lists/Genx sustenance
4 | 3_7 Freeze in Long Lead mode: In Long lead mode, when a record is freezed, the freeze record always appear in a 6X2 format. It is recommended to display and print the exact freezed data as seen on the screen.---Product Team Requirement | Genx12 sustenance |  | Completed | 100 | Mahesh DE | false | Medium |  | 2025-09-04 01:00 |  |  |  |  | Remove Freeze in Long Lead for time being. (As suggested by product Teams) | ¦ Solution: ¦ ¦ --- ¦ ¦ Remove Freeze in Long Lead for time being. <br><br><br><br><br>Ref:<br>Re: Observations from ECG Camp – GenX12i+ Performance and Improvement Suggestions<br>7/03/2025 ¦ |  | Item | sites/ECGTeam761/Lists/Genx sustenance
5 | 2_5 Page Save Mode Issue: In Page Save mode with “ Auto export to USB” enabled, the record is transmitted first followed by preview page which needs to be vice-versa. | Genx12 sustenance |  | Completed | 100 | Mahesh DE | false | Medium |  | 2025-04-09 01:00 | [5 empty] | Sol: ¦ If Auto Export is enabled then there will be no preview.<br><br>If Auto Export is disabled then preview will be shown. <br><br><br><br><br>Ref:<br>Re: Observations from ECG Camp – GenX12i+ Performance and Improvement Suggestions<br>7/03/2025 ¦ ¦ --- ¦ |  | Item | sites/ECGTeam761/Lists/Genx sustenance
6 | 1_4 Auto Export to USB: When “ Auto export to USB” is enabled, and a patient record is exported it does not reflect as “Exported to USB” in Patient Record window. | Genx12 sustenance |  | Completed | 100 | Mahesh DE | false | Medium |  | 2025-04-09 01:00 | 2025-04-04 01:00 | 2025-05-30 01:00 |  |  |  | Ref. Re: Observations from ECG Camp – GenX12i+ Performance and Improvement Suggestions 7/03/2025 |  | Item | sites/ECGTeam761/Lists/Genx sustenance
7 | 6_20 .Add these in measurement matrix (DWconf (%), STslope (deg), Ton (ms), Tdur (ms), T+dur (ms), ) | Genx12 sustenance |  | Completed | 100 | Anubhav Chaudhary;sanjana sampige | false | Medium |  | 2025-04-09 01:00 | [5 empty] | ADD in measurement matrix. DWconf (%), STslope (deg), Ton (ms), Tdur (ms), T+dur (ms), Recommendation from sales Team(GRI Recommendation)Ref Email Subject: Discussion on Recommended ECG Format from Mr. Peter [Glasgow] Dated 1/04/2025 |  | Item | sites/ECGTeam761/Lists/Genx sustenance
8 | 5_19 Detailed report with matrix calculation should be printed in 2pages instead of current practice of 3pages | Genx12 sustenance |  | Dependency (Chk Comment Section) | 100 | Somesh | false | Low |  | 2025-04-09 01:00 | [5 empty] | Recommendation from sales Team(GRI Recommendation) Ref Email Subject: Discussion on Recommended ECG Format from Mr. Peter [Glasgow] Dated 1/04/2025 |  | Item | sites/ECGTeam761/Lists/Genx sustenance
9 | (Product team May 2025) 1_6 .Remove HRV from Interpretation and retain HRV on second page and on LCD. | Genx12 sustenance |  | Completed | 100 | @Mahesh D, Somesh | false | Medium |  | 2025-04-09 | [5 empty] | Recommendation from sales Team(GRI Recommendation) Ref Email Subject: Discussion on Recommended ECG Format from Mr. Peter [Glasgow] Dated 1/04/2025 |  | Item | sites/ECGTeam761/Lists/Genx sustenance
10 | Create software Requirement traceability Matrix | Genx12 sustenance |  | Completed | 100 | Sanjana | false | Medium |  | 2025-04-07 01:00 | [5 empty] | As Genx3 S/W traceability Matrix has Gaps. Discuss with Nitin Sir and Proceed. Ref "06-GENX3-CE\03-DHF\02-SW-DOCS\M_SOFTWARE_REQUIREMENT_TRACEABILITY" |  | Item | sites/ECGTeam761/Lists/Genx sustenance
11 | Create IEC 62304 Check list for Genx12 .Ref F:\04-Genx From Shrikanth\06-GENX3-CE\03-DHF\02-SW-DOCS\N_IEC62304_CHECKLIST | Genx12 sustenance |  | Completed | 100 | Somesh | false | Medium |  | 2025-04-07 01:00 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
12 | Post Market Surveillance and Clinical (For Genx3 ) follow up required after 2021 | Genx3 Sustenance |  | Completed | 100 | Somesh | false | Medium |  | 2025-04-07 01:00 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
13 | Post Market Surveillance required after 2021 | Genx3 Sustenance |  | Completed | 100 | Somesh | false | Medium |  | 2025-04-07 01:00 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
14 | Post Market surveillance Document Required | Genx12 sustenance |  | Completed | 100 | Sanjana | false | Medium |  | 2025-04-07 01:00 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
15 | page Mismatches at all print format at freeze mode- BTR | Genx12 sustenance |  | Completed | 100 | Mahesh DE | false | Medium |  | 2025-04-01 01:00 |  |  |  | 0/2 | interpretation, detailed and off condition- at pages check page number update function;check the print format at freeze mode - Interpretation, Detailed minimal-off | At freeze mode the pages number are mismatched. few print format checked |  | Item | sites/ECGTeam761/Lists/Genx sustenance
16 | PERIODIC_SAMPLES_INCREMENTED - auto export usb mode- periodic duration is 5 , interval is 1 -this condition after complete - the 2 records only transferred | Genx12 sustenance |  | Completed | 100 | Mahesh DE | false | Medium |  | 2025-04-01 01:00 |  |  |  | 1/4 | check the related function;check the "UpdateSampleModes";testing;check the function "UpdateIntervalForPeriodicSample" |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
17 | Verification of Brochure with Technical specification | Genx12 sustenance |  | Completed | 100 | Sanjana | false | Medium |  | 2025-03-25 01:00 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
18 | Display residue is visible on screen while printing-BTR-17 | Genx12 sustenance |  | Completed | 100 | Mahesh DE | false | Medium |  | 2025-03-18 01:00 |  |  |  | 0/1 | check the simulation steps | Auto Mode RT sample>Display and print format : 3x4+3R, Sweep Speed 50mm/sec>Press Freeze>Press START/STOP>Press F3 to print all freeze data>Printout comes in 6x1 format, Freeze display also shows as 6x1, Display residue is visible on screen while printing. Screen showing duration 43-49s, but printout came for 43 seconds only. In case of 120s data, printing came only for 116 seconds |  | Item | sites/ECGTeam761/Lists/Genx sustenance
19 | Debugging GenX12 boards | Genx12 sustenance |  | Completed | 100 | nikhil siby | false | Medium |  | 2025-03-11 01:00 | 2025-03-11 01:00 | 2025-05-06 | nikhil |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
20 | 56 MDD->MDR Manufacturing and test licence QARA | Genx3 Sustenance |  | Not started | 100 |  | false | Medium |  | 2025-02-26 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
21 | 51 MDD->MDR Environmental and reliability test. | Genx3 Sustenance |  | Not started | 100 |  | false | Medium |  | 2025-02-26 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
22 | 49 MDD->MDR Design validation report Chandan | Genx3 Sustenance |  | Not started | 100 |  | false | Medium |  | 2025-02-26 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
23 | 48 MDD->MDR Design validation plan Chandan | Genx3 Sustenance |  | Not started | 100 |  | false | Medium |  | 2025-02-26 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
24 | 47 MDD->MDR Spare part list Somesh | Genx3 Sustenance |  | Completed | 100 | Someshwaran S | false | Medium |  | 2025-02-26 | 2025-03-03 |  |  | 2/3 | Images required for parts and mail send to Mechnaical team-28-02-2025;Update Spare parts list ;Some images is required to complete the list ,as we dont have them . |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
25 | 43 MDD->MDR Technical specifications | Genx3 Sustenance |  | Not started | 100 |  | false | Medium |  | 2025-02-26 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
26 | 42 MDD->MDRproduct training (Nitin Sir) | Genx3 Sustenance |  | Not started | 100 |  | false | Medium |  | 2025-02-26 | [5 empty] | LL Mode training Doc need to be created. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
27 | IEC60601-1-2 test report. | GenX3 docs - MDD to MDR conversion |  | Not started | 100 |  | false | Medium |  | 2025-02-26 |  |  |  | 0/1 | After completing complisnce testing on SMPS of Genx3 report has to be updated |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
28 | Convert RT view Documents into New Format | No Bucket |  | Dependency (Chk Comment Section) | 100 | sanjana sampige | false | Medium |  | 2025-02-14 |  |  |  | 1/1 | Email the list of documents |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
29 | Initiate email for bridge PCB connector failure (Root Cause) | Genx12 sustenance |  | Dependency (Chk Comment Section) | 100 |  | false | Medium |  | 2025-02-14 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
30 | if freeze screen showing 4/4- the NEXT PAGE SHOULD BE BLOCKED | Genx12 sustenance |  | Completed | 100 | Mahesh DE | false | Medium |  | 2025-02-13 | [5 empty] | freeze screen showing below pages- number NUMBER Display Menu 1/4- NEXT.PAGES , 2/4- PREV-PAGES and NEXT.PAGES ,3/4- PREV PAGES and NEXT.PAGES ,4/4- PREV.PAGES |  | Item | sites/ECGTeam761/Lists/Genx sustenance
31 | Create Revision History for Dot Net Board Version | Genx12 Design Transfer |  | HOLD(Dependency) | 100 | Anubhav Chaudhary | false | Medium |  | 2025-01-28 | [5 empty] | 1.5 -> Genx3 1.6 -> Not modified( Wifi testing) 1.7 -> Not modified(B-T testing) 1.8 -> CAD FOUNDATION 1.9 -> Genx12 Release. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
32 | Disp format changes in one scenario | Genx12 sustenance |  | Completed | 100 | Mahesh DE | false | Medium |  | 2025-01-16 | 2025-02-06 |  |  | 5/6 | after completing acquisition - freeze screen comes on the Display- then again press START/STOP butto;check the Display format and printout;if display 6x2+1 after escaping home screen comes to the - 6x1 ;if display 3x4+1 after escaping home screen comes to the - 3x4;after Escaping - main Home screen comes to the 6x1;check the simulation step - page save -RT sample-- Display format 3x4+3 - press START/STOP button | Set PGSV Screen (or Freeze Mode)->Disp 3x4+3R(Scroll View) ----> Print Rec -----> After print Return Home Screen , Disp Format will Change. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
33 | 2. In certain instances, the HR displayed on the screen differs from the HR printed. GENERAL OBSERVATIONS: Negative | Genx12 sustenance |  | Completed | 100 |  | false | Medium |  | 2024-09-07 01:00 |  |  |  | 0/1 | Internal Discussion Required | EXPLAINED:::: HR computed on screen is real time HR whereas HR printed on paper comes from 10 sec ecg data provided to Glasgow Algorithm. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
34 | 2.4 At times occasionally, even when all the leads are connected, the waveform is not displayed in any format GenX12i+ – YKJA4G1002 | Genx12 sustenance |  | Completed | 100 | Anubhav Chaudhary;chandan kumar | false | Medium |  | 2024-09-07 01:00 |  |  |  | 1/2 | Simulation required;Discussion needed | - As per our understanding it will only happen when any of the limb lead is disconnected or if patient cable is loose. - Not observed during our testing. Simulation Required for identifying the root cause. - Need V&V help to simulate. Please help us with simulation steps. Simulation will help us to fix it quickly. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
35 | 1.4 Without entering any patient details, symbols, special characters, and "Sex: D" appeared in the patient record. These records were also not transmitted via USB. GenX12i –- XQJA4G1002 | Genx12 sustenance |  | Completed | 100 | Anubhav Chaudhary;chandan kumar | false | Medium |  | 2024-09-07 01:00 |  |  |  | 0/1 | Need Simulation | - Not observed during our testing. Simulation Required for identifying the root cause. - We are trying to simulate this. If steps are provided it will be useful Need V&V help to simulate. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
36 | BARCODE SCANNER feasibility. | Genx12 sustenance |  | Not started | 100 |  | false | Medium |  | 2024-07-29 01:00 |  |  |  | 0/4 | Check and calculate the time for full scan and one complete workflow.;Understand and check feasibility against our system.;Discuss and understand with product team for work flow.;Requirement document required from product team | Will be done on New platform |  | Item | sites/ECGTeam761/Lists/Genx sustenance
37 | LCD Brightness shift to PWM control with GUI option addition. | Feature Addition |  | HOLD(Dependency) | 100 | Someshwaran S | false | Medium |  | 2024-07-23 01:00 |  |  |  | 6/8 | Touch enable for gui option;PWM enable -8 inch;Merge changes with original code.;Verifing the code with anubhav sir ;Add Gui option-8 inch;Add Gui option;PWM enable;Asses the code for implementation | Will be completed when new hardware(Main Board upgraded) will be released. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
38 | Explore overlapping with trace shift with max and min on plot. | Genx12 sustenance |  | Not started | 100 |  | false | Low |  | 2024-07-23 01:00 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
39 | Genx3 4.3 Display cost productivity activity | Genx3 Sustenance |  | Completed | 100 |  | false | Medium |  | 2024-07-22 01:00 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
40 | Long XE SMPS for Genx1, Genx3 | Cost Productivity |  | Dependency (Chk Comment Section) | 100 | nikhil siby | false | Medium |  | 2024-07-22 01:00 | 2024-08-26 01:00 |  |  | 6/7 | BOM update;Pre-Compliance Testing at Jigani BPLMTPL Factory.;Creating Report(Observation and Conclusion).;Identify and implementing changes to make compatible with existing M/C;Functional testing.;Dry testing;Line and Load Regulation testing. | Dependencies:- Will depend on ECN for change. Aman (Pre-compliance) Somesh and Nitin (BOM)(CCL) GenX12 & GenX3 testing completed GenX1 pending |  | Item | sites/ECGTeam761/Lists/Genx sustenance
41 | Dot Net Board Transfer Speed increase assessment(Break up time check SPI transfer and PDF creation time) | To do |  | Completed | 100 |  | false | Low |  | 2024-07-22 01:00 | [5 empty] | Dot Net Board terminal window uart is not working on blocking mode so blocking it will not have any impact on speed. Dot Net Board commands transfer at 115200 Baud Rate. SPI Clock rate is 4 MHZ |  | Item | sites/ECGTeam761/Lists/Genx sustenance
42 | VCG Implementation to check with GRI and check expectation with marketing team. | Feature Addition |  | Dependency (Chk Comment Section) | 100 |  | false | Medium |  | 2024-07-09 01:00 |  |  |  | 1/20 | Create Design Feasibility Report (Dependency Sample Report from Inder and Nitin Sir);Refine SRS as per feedback and Risk assessment;Create and execute Home Screen with VCG for all three plane plots.;Review SRS with ECG team, V& V and Product Team;Convert MRD to High Level requirement;Create High level design for VCG home screen display in paint.;POC:::::JUST IMPLEMENT and execute VCG on platform.(result to be checked through debugger or UART) ;MRD Received .;Risk Assessment to VCG feature;Add all key(keypad) functionality on Home Screen for VCG.;Test all functionality on home screen.;Create all touch functionality on home screen for VCG. ;Execute memory storage for VCG data.;Start implementing the VCG with separate Mode creation.;Draw Print format image in paint to get requirement.;Execute Live VCG Print.;Create memory storage format for VCG data.;Review of Proof of concept with all stake holders.;Create SRS for VCG;Plan Print for VCG with display.(Print from RAM buffer) | Will be part of new platform. 21. ECG record storage retrieval for Display.(First chk in Paint) 22. Implement the Record print same as live print (with REC COPY) printed. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
43 | 4.3 Marker ON- but missed printing twice and left a blank print on patient connect Genx12 – Gujarat (Ahmedabad) | GenX12i Series observations from Camp |  | Completed | 100 | Anubhav Chaudhary;chandan kumar | false | Medium | 2025-04-10 11:30 | 2024-09-07 01:00 |  |  | chandan kumar | 0/2 | Simulation;Waiting for Machine | - Not observed during our testing. Simulation Required for identifying the root cause. - Need V&V help to simulate. We are trying to simulate this. If steps for simulation are provided it will be useful. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
44 | 3_15 Remove R’’ and S’’ amplitude and duration in Measurement matrix. | Genx12 sustenance |  | Completed | 100 | Anubhav Chaudhary;sanjana sampige | false | Medium | 2025-04-09 11:30 | 2025-04-09 01:00 |  |  | sanjana sampige |  |  | Recommendation from sales Team(GRI Recommendation) Ref Email Subject: Discussion on Recommended ECG Format from Mr. Peter [Glasgow] Dated 1/04/2025 |  | Item | sites/ECGTeam761/Lists/Genx sustenance
45 | 7_21. Add “notch” to “QT int” or insert values for notch into this row. Notch count should be 0, 1, or maybe 2.We are writing Duration in Notch field. | Genx12 sustenance |  | Completed | 100 | Anubhav Chaudhary;sanjana sampige | false | Medium | 2025-04-09 11:30 | 2025-04-09 01:00 |  |  | sanjana sampige |  |  | Recommendation from sales Team(GRI Recommendation) Ref Email Subject: Discussion on Recommended ECG Format from Mr. Peter [Glasgow] Dated 1/04/2025 |  | Item | sites/ECGTeam761/Lists/Genx sustenance
46 | 36 MDD->MDR Risk management report is different - one is excel doc and the other is a word doc | Genx3 Sustenance |  | Completed | 100 |  | false | Medium | 2025-04-08 11:30 | 2025-02-26 |  |  | Anubhav Chaudhary |  |  | Will be done MEDITIX |  | Item | sites/ECGTeam761/Lists/Genx sustenance
47 | 34 MDD->MDR Risk benefit analysis formats are different | Genx3 Sustenance |  | Completed | 100 |  | false | Medium | 2025-04-08 11:30 | 2025-02-26 |  |  | Anubhav Chaudhary |  |  | Will be done MEDITIX |  | Item | sites/ECGTeam761/Lists/Genx sustenance
48 | 35 MDD->MDR Risk management plan format is different | Genx3 Sustenance |  | Completed | 100 |  | false | Medium | 2025-04-08 11:30 | 2025-02-26 |  |  | Anubhav Chaudhary |  |  | Will be done MEDITIX |  | Item | sites/ECGTeam761/Lists/Genx sustenance
49 | 33 MDD->MDR Annexure documents not present in Risk benefit analysis in GenX3 | Genx3 Sustenance |  | Completed | 100 |  | false | Medium | 2025-04-08 11:30 | 2025-02-26 |  |  | Anubhav Chaudhary |  |  | Will be done MEDITIX |  | Item | sites/ECGTeam761/Lists/Genx sustenance
50 | MDD->MDR = 6. Software Unit Verification Report share with QARA | Genx3 Sustenance |  | Completed | 100 |  | false | Medium | 2025-04-08 11:30 | 2025-02-02 |  |  | Anubhav Chaudhary | 0/2 | Software unit implementation;DONE |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
51 | MDD->MDR 5. Share USVR location with QARA | Genx3 Sustenance |  | Completed | 100 |  | false | Medium | 2025-04-08 11:30 | 2025-02-04 |  |  | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
52 | Aligning Doc Gaps for Genx12 | AUDIT APRIL 2025 |  | Completed | 100 | Someshwaran S | false | Medium | 2025-04-06 11:30 | 2025-04-02 01:00 |  |  | Someshwaran S |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
53 | Aligning document Gaps for Genx3 | AUDIT APRIL 2025 |  | Completed | 100 | Someshwaran S | false | Medium | 2025-04-06 11:30 | 2025-04-02 01:00 |  | 2025-04-04 01:00 | Someshwaran S |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
54 | Brochure comparison | AUDIT APRIL 2025 |  | Completed | 100 | sanjana sampige | false | Medium | 2025-04-06 11:30 | 2025-04-02 01:00 |  |  | sanjana sampige |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
55 | Close the Genx3 NC Related to SRN and SAN .Update (Replace0 SAN<-> SRN in DOF Document.) | Genx3 Sustenance |  | Completed | 100 | Someshwaran S | false | Medium | 2025-03-31 11:30 | 2025-03-28 01:00 |  |  | Someshwaran S |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
56 | CCL Validation(Certificate) for Genx12 | Genx12 sustenance |  | Completed | 100 | Someshwaran S | false | Medium | 2025-03-30 11:30 | 2025-03-25 01:00 |  | 2025-03-27 01:00 | Someshwaran S |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
57 | CCL validation for Gen3 | Genx3 Sustenance |  | Completed | 100 | Someshwaran S | false | Medium | 2025-03-30 11:30 | 2025-03-25 01:00 |  | 2025-03-27 01:00 | Someshwaran S |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
58 | Creating Response(Product Team) for Peter Mc Farlane(GRI) | Genx3 Sustenance |  | Completed | 100 | sanjana sampige | false | Medium | 2025-03-26 11:30 | 2025-03-25 01:00 |  | 2025-03-25 01:00 | sanjana sampige |  |  | Checking feasibility DW Config(%) ST Slope (Deg) T+ dur (ms) |  | Item | sites/ECGTeam761/Lists/Genx sustenance
59 | Brochure Comparison for GenX3 and GenX12 | No Bucket |  | Completed | 100 | sanjana sampige | false | Medium | 2025-03-25 11:30 | 2025-03-26 01:00 |  |  | sanjana sampige |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
60 | Co-ordination and creating Documents for QARA Regarding BIS Document Submission | Genx3 Sustenance |  | Completed | 100 | sanjana sampige | false | Medium | 2025-03-25 11:30 | 2025-03-25 01:00 |  | 2025-03-28 01:00 | sanjana sampige |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
61 | LCD cable drawings - update according to the colour mentioned (VV12) | Vivid Vue |  | Completed | 100 | samiksha shetty | false | Medium | 2025-03-25 11:30 | 2025-03-26 01:00 |  |  | samiksha shetty | 4/4 | 414-E-060(MITSHUBISHI 10_12);412-E-042(MITSHUBISHI 8);414-E-040(AUO10.4 SHIELD);414-E-039(AUO12.1SHIELD) |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
62 | Display issue - reworks in tech board (VV15) | Vivid Vue |  | Completed | 100 | samiksha shetty | false | Medium | 2025-03-25 11:30 | 2025-03-26 01:00 |  |  | samiksha shetty |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
63 | Split cable verification - 50 PIN to 20 & 30 PIN (VV15) | Vivid Vue |  | Completed | 100 | samiksha shetty | false | Medium | 2025-03-25 11:30 | 2025-03-26 01:00 |  |  | samiksha shetty |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
64 | LongXC power cables length verification - 5 Qty (VV15) | Vivid Vue |  | Completed | 100 | samiksha shetty | false | Medium | 2025-03-25 11:30 | 2025-03-26 01:00 |  |  | samiksha shetty |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
65 | New LCD display cable verification (VV15) | Vivid Vue |  | Completed | 100 | samiksha shetty | false | Medium | 2025-03-25 11:30 | 2025-03-26 01:00 |  |  | samiksha shetty |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
66 | Verifying the quantities of the components received for assembling 3 units (VV15) | Vivid Vue |  | Completed | 100 | samiksha shetty | false | Medium | 2025-03-25 11:30 | 2025-03-07 | 2025-03-06 |  | samiksha shetty | 3/3 | Kepad board; ITC board;Tech board | 11/03/2025 - received new set of components |  | Item | sites/ECGTeam761/Lists/Genx sustenance
67 | Genx3i Folder Creation for Audit Purpose | Genx3 Sustenance |  | Completed | 100 | Someshwaran S | false | Medium | 2025-03-25 11:30 | 2025-03-25 01:00 | 2025-03-25 01:00 | 2025-03-25 01:00 | Someshwaran S |  |  | Created genx3i folders with old documents |  | Item | sites/ECGTeam761/Lists/Genx sustenance
68 | Battery PCB male connector -veriifcation | Genx12 sustenance |  | Completed | 100 | Someshwaran S | false | Medium | 2025-03-25 11:30 | 2025-03-18 01:00 | 2025-03-18 01:00 | 2025-03-25 01:00 | Someshwaran S |  |  | Received 1 no of male connector from kaynes for battery charger pcb. Ask suganthi mam to mount on the machine. Need to check the mating with main pcb-kaynes Checked wit kaynes,kaynes need to comeback |  | Item | sites/ECGTeam761/Lists/Genx sustenance
69 | Genx3 Folder Creation for Audit purpose | Genx3 Sustenance |  | Completed | 100 | Someshwaran S | false | Medium | 2025-03-24 11:30 | 2025-03-25 01:00 | 2025-03-24 01:00 | 2025-03-24 01:00 | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
70 | Neonatal cable test | Genx12 Design Transfer |  | Completed | 100 | beryl paulson | false | Medium | 2025-03-23 11:30 | 2025-01-13 |  |  | Nitin Chavan | 3/4 | 1) Functional testing;3) Defib test at jigani;report making (Dependency on anil(PED));2)Compared print with existing patient cable | On Hold as dependency on ANIL(Q/C) 1)Report making in progress 2) need to conduct test in GenX3 unit 3)Already send Cables to factory to test on GenX3 machine. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
71 | Give two machines to V&V for SMPS testing.(LongXc and TDK lamda) | Genx12 Design Transfer |  | Completed | 100 | nikhil siby;chandan kumar | false | Medium | 2025-03-23 11:30 | 2025-03-03 |  |  | Nitin Chavan |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
72 | Encoder Reliability | Vivid Vue |  | Completed | 100 | nikhil siby | false | Medium | 2025-03-23 11:30 | 2025-03-03 | 2025-03-03 | 2025-03-31 01:00 | nikhil siby | 5/5 | Report ;5000/360000 turns clockwise (03-03-2025);11040/36500 cycles anticlockwise (07-03-2025);18250/36500 cycles anticlockwise (10-03-2025);17000/360000 turns clockwise (04-03-2025) | 17000/360000 turns clockwise (04-03-2025) Itc board got damaged and proceeding with the rework(05-03-2025) Display is also damaged so changed the display(06-03-2025) 18250 cycles anticlockwise (10-03-2025) 36500 cycles completed (18250 CLOCKWISE & 18250 ANTICLOCKWISE) (11-03-2025) |  | Item | sites/ECGTeam761/Lists/Genx sustenance
73 | Notch filter testing | Genx12 sustenance |  | Completed | 100 | Someshwaran S;sanjana sampige | false | Medium | 2025-03-20 11:30 | 2025-03-21 01:00 | 2025-03-20 01:00 | 2025-03-21 01:00 | Someshwaran S |  |  | Test with Function generator by giving frequencies from 10hz to 100 hz in Genx12 in AC filter off,50hz,60hz.Compare the same with 9108,GE Machine.-Somesh Test with Simulator by giving ECG signal and conduction signal by 50hz,60hz muscular in Genx12,9108,GE Machines -Sanajna |  | Item | sites/ECGTeam761/Lists/Genx sustenance
74 | Basic Info form send to Astute along with CCL | Genx12 Design Transfer |  | Completed | 100 | Anubhav Chaudhary;Someshwaran S;sanjana sampige | false | Medium | 2025-03-20 11:30 | 2025-03-21 01:00 | 2025-03-21 01:00 | 2025-03-21 01:00 | Someshwaran S |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
75 | Simulation for GenX12 Bugs | No Bucket |  | Completed | 100 | sanjana sampige | false | Medium | 2025-03-20 11:30 | 2025-03-17 01:00 |  |  | sanjana sampige |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
76 | GENX12 Software Maintenance plan | Genx12 sustenance |  | Completed | 100 | Someshwaran S | false | Medium | 2025-03-19 11:30 | 2025-03-20 01:00 | 2025-03-20 01:00 | 2025-03-20 01:00 | Someshwaran S |  |  | Created Software Maintenace plan and Software Configuration Management document |  | Item | sites/ECGTeam761/Lists/Genx sustenance
77 | Display check for the new set of tech boards after rework (VV15) | Vivid Vue |  | Completed | 100 | samiksha shetty | false | Medium | 2025-03-19 11:30 | 2025-03-20 01:00 |  |  | samiksha shetty |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
78 | Continuity check - Tech board inductors (VV15) | Vivid Vue |  | Completed | 100 | samiksha shetty | false | Medium | 2025-03-19 11:30 | 2025-03-13 01:00 | 2025-03-13 01:00 |  | samiksha shetty | 2/2 | Check for all the received boards.;Rework to be done for the incorrect boards. |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
79 | Paper Stop Cutting Edge(BTR) | Genx12 sustenance |  | Completed | 100 | Mahesh DE | false | Medium | 2025-03-19 11:30 | 2025-01-16 |  |  | Mahesh DE |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
80 | Chandan Regressing testing non simulation issue closure and fixed issue report to shar e with Chandan. | Genx12 sustenance |  | Completed | 100 | Mahesh DE | false | Medium | 2025-03-19 11:30 | 2025-02-03 |  |  | Mahesh DE |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
81 | Wiring diagram - check & update (VV15) | Vivid Vue |  | Completed | 100 | samiksha shetty | false | Medium | 2025-03-18 11:30 | 2025-03-11 01:00 | 2025-03-10 01:00 |  | samiksha shetty |  |  | Check done. Correcting the mistakes found. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
82 | Update GenX3 User manual | No Bucket |  | Completed | 100 | Someshwaran S;sanjana sampige | false | Medium | 2025-03-18 11:30 | 2025-03-17 01:00 |  |  | sanjana sampige |  |  | Compared and added items from the user manual edited and commented by Nitin sir |  | Item | sites/ECGTeam761/Lists/Genx sustenance
83 | Manufacturing details list for genx3 ccl | Genx3 Sustenance |  | Completed | 100 | Someshwaran S | false | Medium | 2025-03-17 11:30 | 2025-03-18 01:00 |  | 2025-03-18 01:00 | Someshwaran S |  |  | Added manufacturing details with address and certificate &test report numbers for genx3. Removed FSP Smps and added Longxc,arteysn smps Removed BPL Patient cable from the list |  | Item | sites/ECGTeam761/Lists/Genx sustenance
84 | Verify Cables which came from kaynes | Genx12 sustenance |  | Completed | 100 | Someshwaran S | false | Medium | 2025-03-17 11:30 | 2025-03-18 01:00 | 2025-03-18 01:00 | 2025-03-18 01:00 | Someshwaran S |  |  | Received ground cables from kaynes after rework with 3.7mm ID.(439-E-038,039,040) Assembled on machine and shown to nitin sir LCD Cable is received-3 no's. But it is not as per drawings(twisted pairs are not maintained properly) Created report for verification |  | Item | sites/ECGTeam761/Lists/Genx sustenance
85 | Inspection procedure Vivid vue 15 | Vivid Vue |  | Completed | 100 | beryl paulson | false | Medium | 2025-03-16 11:30 | 2025-01-20 | 2025-02-05 |  | beryl paulson | 6/6 | Given to Mechanical team for review of mechanical inspection 12/2/2025;Recived feedback from mechanical 24/2/2025;Ready for 1st review.14/2/2025;Completed from my side with DATA from User manual and given for Review 14/02/2025;1st Review done 26/2/2025;Fetching required labels and drawings from mechanical | 1)Given to review for mechanical team for the section mechanical inspection12/2/2025 2)Completed from my side with DATA from User manual and given for Review 14/02/2025 3)Ready for 1st review.14/2/2025 4)Recived feedback from mechanical 24/2/2025 4) 1st Review done 26/2/2025 5)Send mail to software team 4/3/2025 6)recieved comments from software team4/3/2025 7)As per Nitin sir comment given to sindhu for further steps on 10/3/2025 |  | Item | sites/ECGTeam761/Lists/Genx sustenance
86 | Prepare one unit for user manual inspection | Vivid Vue |  | Completed | 100 | nikhil siby | false | Medium | 2025-03-16 11:30 | 2025-03-17 01:00 |  |  | nikhil siby |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
87 | 45 MDD->MDR User manual Sindhu | Genx3 Sustenance |  | Completed | 100 | Sindhu Kashyap | false | Medium | 2025-03-12 11:30 | 2025-02-26 | 2025-03-05 | 2025-03-12 01:00 | Sindhu Kashyap |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
88 | Display cable rework to be done (VV8) | Vivid Vue |  | Completed | 100 | samiksha shetty | false | Medium | 2025-03-11 11:30 | 2025-03-12 01:00 |  | 2025-03-12 01:00 | samiksha shetty |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
89 | Genx3 user manual changes (Medtronics' inputs) | Genx3 Sustenance |  | Completed | 100 | samiksha shetty | false | Medium | 2025-03-11 11:30 | 2025-03-12 01:00 |  |  | samiksha shetty |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
90 | 37 MDD->MDR Design quality plan format different | Genx3 Sustenance |  | Completed | 100 |  | false | Medium | 2025-03-11 11:30 | 2025-02-26 |  |  | sanjana sampige |  |  | No need for format update |  | Item | sites/ECGTeam761/Lists/Genx sustenance
91 | 39 MDD->MDR Product life declaration format is different | Genx3 Sustenance |  | Completed | 100 |  | false | Medium | 2025-03-11 11:30 | 2025-02-26 |  |  | sanjana sampige |  |  | No need for format update |  | Item | sites/ECGTeam761/Lists/Genx sustenance
92 | 52 MDD->MDR Product brochure | Genx3 Sustenance |  | Completed | 100 |  | false | Medium | 2025-03-11 11:30 | 2025-02-26 |  |  | sanjana sampige |  |  | No need for update |  | Item | sites/ECGTeam761/Lists/Genx sustenance
93 | 55 MDD->MDR Bio compatibility test report Done | Genx3 Sustenance |  | Completed | 100 |  | false | Medium | 2025-03-11 11:30 | 2025-02-26 |  |  | sanjana sampige |  |  | No need for format update. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
94 | 57 MDD->MDR Report of clinical evaluation Done | Genx3 Sustenance |  | Completed | 100 |  | false | Medium | 2025-03-11 11:30 | 2025-02-26 |  |  | sanjana sampige |  |  | No need for format update |  | Item | sites/ECGTeam761/Lists/Genx sustenance
95 | COPY functionality enlist in SRS,SVP. | Genx12 Design Transfer |  | Completed | 100 |  | false | Medium | 2025-03-11 11:30 | 2024-09-11 01:00 |  |  | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
96 | Add Dot Net Board Software in Genx3 Horizontal impact. | Genx12 Design Transfer |  | Completed | 100 |  | false | Medium | 2025-03-11 11:30 | 2025-01-21 |  |  | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
97 | 29 MDD->MDR Inspection procedure is different | Genx3 Sustenance |  | Completed | 100 |  | false | Medium | 2025-03-11 11:30 | 2025-02-26 |  |  | sanjana sampige | 0/1 | DONE By Samiksha |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
98 | 27 MDD->MDR GenX12 has SRN and axf files whereas GenX3 has SAN | Genx3 Sustenance |  | Completed | 100 |  | false | Medium | 2025-03-11 11:30 | 2025-02-26 |  |  | sanjana sampige | 0/1 | DONE |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
99 | 24 MDD->MDR SRS file not present in GenX from srikanth | Genx3 Sustenance |  | Completed | 100 |  | false | Medium | 2025-03-11 11:30 | 2025-02-26 |  |  | sanjana sampige | 0/1 | DONE |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
100 | Test Effected bridge PCB with 1 mbps baud rate | Genx12 sustenance |  | Completed | 100 | Anubhav Chaudhary | false | Medium | 2025-03-11 11:30 | 2025-02-21 | 2025-03-10 01:00 | 2025-03-10 01:00 | Anubhav Chaudhary |  |  | WE have tested it and there is no issue on 921600 BR from bridge PCB side. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
101 | Genx12 Keyboard version read | Genx12 sustenance |  | Completed | 100 | Anubhav Chaudhary | false | Medium | 2025-03-11 11:30 | 2025-02-25 | 2025-03-11 01:00 | 2025-03-12 01:00 | Anubhav Chaudhary |  |  | KEYBOARD_VERSION_READ_FIX Added this on 11/03/2025 As keyboard version is only required in System info menu, so we are reading Keyboard version when system info is pressed. To avoid any confusion we read version 5 times(with in 1Sec). |  | Item | sites/ECGTeam761/Lists/Genx sustenance
102 | Test Genx12 board with bridge PCB with increased Baud rate | Genx12 sustenance |  | Completed | 100 | Anubhav Chaudhary | false | Medium | 2025-03-11 11:30 | 2025-02-25 | 2025-02-25 | 2025-03-07 | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
103 | Genx12 and Genx12+ : Initiate and drive activity for Compliance test(Astute Lab) for LongXc and TDK SMPS. | Genx12 sustenance |  | Completed | 100 | nikhil siby | false | Medium | 2025-03-10 11:30 | 2025-02-25 | 2025-02-28 | 2025-03-11 01:00 | nikhil siby | 3/3 | Two SMPS TDK Lamda need to be send(11/03/2025);Basic Testing for two units;M/C assembly | 1. One device is ready without dot net board (28-02-2025) 2. 2nd device need ferrite core. Dot net board and ferrite core requested to factory, once the items received machines can be assemble (28-02-2025) 3. Got 2 Dot net boards(07-03-2025) Both devices are ready (07-03-2025) |  | Item | sites/ECGTeam761/Lists/Genx sustenance
104 | Vivid Vue Battery testing for 2A | Vivid Vue |  | Completed | 100 | beryl paulson | false | Medium | 2025-03-09 11:30 | 2025-02-11 | 2025-02-11 | 2025-03-11 01:00 | beryl paulson | 7/7 | Battery 3 discharging;Charging done 2A battery 2;Discharging in progress;Battery 3 charging;charging time analysis for 2A battery;Discharged 2nd battery; rework for charging current 2A | 1)Rework done for charging current loc:R21 - 13.79V (11/02/2025) 2)To initiating charging time analysis ,discharging the battery (12/02/2025) 3)Charging done for 2A charging current(battery 1)-12/2/2025 4)Discharge completed for 2A charging current(battery 1)-13-2-2025 5)Issue in ITC board found (battery 2)-14-02-2025 6)Did Rework for 3.3 votage since display was not getting on, shorting issue saw in R581 (0E) resistor(14/2/2025) Debug Did for 3.3V 1.L31 was removed and checked for short andkept it back 2.R580 removed and checked for short and not connected 3.R581 removed and checked for short,(shorting was there) ,not connected 4.removed C87,C88,C89,C90,C91 was removed and checked for short and not connected Issue:(Display was not getting on) voltage at pin 1,2-12v not coming voltage at pin 7-3.3V not coming 7)Cant able to do discharge and charging due to issue 8)Shared excel sheet with patient monitor team.10/3/2025 Battery analysis: Battery 1: charging time: 3Hour 14 min Discharging time: 4 Hours Battery 2: charging time: 3 hours Discharging time: 4hour 13 min Battery 3: charging time:3 Hours Discharging time: 4 Hour 2 min |  | Item | sites/ECGTeam761/Lists/Genx sustenance
105 | Vivid vue 15 battery testing (1A) | Vivid Vue |  | Completed | 100 | beryl paulson | false | Medium | 2025-03-09 11:30 | 2025-02-28 | 2025-02-27 | 2025-03-10 01:00 | beryl paulson | 12/12 | Battery 3 Discharging 3/3/025-started ,3/3/3035-end;battery 2 charging 28/2/2025-started, 3/3/2025-end;battery 6 Charging-6/3/2025-started.7/3/2025-End;battery 1 charging 27/2/2025-started ,28/2/2025-end;Battery 3 Charging -3/3/2025-started,4/3/2025-end;battery 4 Discharging-4/3/2025-started ,4/3/2025-end;battery 5 Discharging-6/3/2025-started,6/3/2025-End;battery 1 discharging 27/2/2025;battery 2 discharging 28/2/2025-Started, 28/2/2025-end;battery 5 Charging 5/3/2025-started,5/3/2025-end;battery 6 Discharging-6/3/2025-started,6/3/2025-end;battery 4 Charging-4/3/2025-started,4/3/2025-end | Discussion points: 1)As per the discusion 1A battery(3.18 Ah)part no: 3S1P5SWB-01-R1 started again for charging and discharge (26/2/2025) 2)altogether 6 battery (1A) is present 3)Shared excel sheet of battery analysis with patient monitor team.10/3/2025 charging Analysis Battery 1 (1A)-Discharging time-2 hour 3 min Charging time-2Hour 25 min Battery 2 (1A) Discharging time-2hour Charging time-2Hour 26 min Battery 3 (1A) Discharging time-1Hour 58 min Charging time-2 hour 54 min Battery 4 (1A) Discharging time-2Hour 3min charging time-2hour 46 min Battery 5 (1A) Discharging time- 1Hour 50 min Charging time- 3Hour 20 min Battery 6 (1A) Discharging time- 1Hour 45 min Charging time- 2Hour 38 min |  | Item | sites/ECGTeam761/Lists/Genx sustenance
106 | Make one GenX12 unit 9 inch for chandan | Genx12 sustenance |  | Completed | 100 | nikhil siby | false | Medium | 2025-03-09 11:30 | 2025-03-07 |  | 2025-03-10 01:00 | nikhil siby |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
107 | 22 MDD->MDR Schematic diagram approval note not present for GenX3 | Genx3 Sustenance |  | Completed | 100 | Someshwaran S | false | Medium | 2025-03-06 10:30 | 2025-02-26 | 2025-03-06 | 2025-03-06 | Someshwaran S |  |  | Created Schematic approval notes for the schematics |  | Item | sites/ECGTeam761/Lists/Genx sustenance
108 | 41 MDD->MDRDesign Part List (Genx3) Somesh | Genx3 Sustenance |  | Completed | 100 | Someshwaran S | false | Medium | 2025-03-06 10:30 | 2025-02-26 | 2025-03-05 | 2025-03-07 | Someshwaran S |  |  | Updated the comments which came from EMS Micron Linked the GENX3 CE Related parts(Cushion ,carton,ECG cable ,core,shield ) in SAP |  | Item | sites/ECGTeam761/Lists/Genx sustenance
109 | 50 MDD->MDR Critical component list Somesh | Genx3 Sustenance |  | Completed | 100 | Someshwaran S | false | Medium | 2025-03-06 10:30 | 2025-02-26 | 2025-03-02 | 2025-03-06 | Someshwaran S | 2/2 | Update CCL;review form nitin sir | Added Longxc ,Emerson Smps and neonatal cables updated PCB revisions Expecting Emerson smps standards and files |  | Item | sites/ECGTeam761/Lists/Genx sustenance
110 | 75. MDD->MDR Create One DRR for Genx3 (Final One) | Genx3 Sustenance |  | Completed | 100 | Someshwaran S | false | Medium | 2025-03-06 10:30 | 2025-03-03 |  |  | Someshwaran S |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
111 | Update alternate SMPS & neonatal cablein Critical Part list for Genx3 and Genx12. | Genx12 sustenance |  | Completed | 100 | Someshwaran S | false | Medium | 2025-03-06 10:30 | 2025-02-26 | 2025-02-26 |  | Someshwaran S | 5/5 | Adding in CCL and putting necessary documents in CCL folder; LongXC and TDK lambada SMPS for Genx12 -Somesh;Getting price and Lead time from SCM ;include longxc and emerson SMPS for Genx3 -Nikhil;Creating SAP code |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
112 | Vivid vue relaibility test for Touch | Vivid Vue |  | Completed | 100 | Sindhu Kashyap | false | Medium | 2025-03-06 10:30 | 2025-02-05 | 2025-02-07 | 2025-03-07 | nikhil siby | 14/14 | 21000/73000 touches completed (14-02-2025) (3000);11500/73000 touches completed (10-02-2025);14500/73000 touches completed (11-02-2025);35000/73000 touches completed (06-03-2025);4500/73000 touches completed;30000/73000 touches completed (04-03-2025);9000/73000 touches completed (07-02-2025);23500/73000 touches completed (17-02-2025) (2500);6500/73000 touches completed;27000/73000 touches completed (03-03-2025);18000/73000 touches completed (12-02-2025) 2500 touches;15500/73000 touches completed (12-02-2025);24000/73000 touches completed (18-02-2025) (500);33000/73000 touches completed (05-03-2025) | 35000/73000 touches completed (06-03-2025) Completed the test with 35000 touches. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
113 | Purchase BOM excel to be filled (VV15) | Vivid Vue |  | Completed | 100 | samiksha shetty | false | Medium | 2025-03-06 10:30 | 2025-03-05 | 2025-03-05 |  | samiksha shetty | 5/5 | VIVIDVUE15 MASIMO MS AGM KIT;VIVIDVUE15 MASIMO SS AGM WITHOUT O2 KIT;VIVIDVUE15 MASIMO SS AGM WITH O2 KIT;VIVIDVUE15 MASIMO MS ETCO2 KIT;VIVIDVUE15 MASIMO SS ETCO2 KIT |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
114 | BIS audit related reports for GenX12 | No Bucket |  | Completed | 100 | samiksha shetty | false | Medium | 2025-03-06 10:30 | 2025-03-05 |  | 2025-03-07 | samiksha shetty | 1/1 | Raw materials details with the required BIS certificates & test reports submitted to QARA |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
115 | Update user manual changes in RMF | No Bucket |  | Completed | 100 | sanjana sampige | false | Medium | 2025-03-06 10:30 | 2025-03-07 |  |  | sanjana sampige |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
116 | 61. MDD->MDR Genx3 Product Release Note | Genx3 Sustenance |  | Completed | 100 | sanjana sampige | false | Medium | 2025-03-06 10:30 | 2025-03-03 |  |  | sanjana sampige |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
117 | Update user manuals with cleaning information | No Bucket |  | Completed | 100 | sanjana sampige | false | Medium | 2025-03-05 10:30 | 2025-03-06 |  |  | sanjana sampige |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
118 | 23 MDD->MDR Gerber approval note | Genx3 Sustenance |  | Completed | 100 | nikhil siby | false | Medium | 2025-03-04 10:30 | 2025-02-26 | 2025-03-05 |  | nikhil siby |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
119 | 44 MDD->MDR Board alignment Nikhil | Genx3 Sustenance |  | Completed | 100 | sanjana sampige | false | Medium | 2025-03-04 10:30 | 2025-02-26 |  |  | sanjana sampige |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
120 | 62. MDD->MDR Design Input Requirement formats are different | Genx3 Sustenance |  | Completed | 100 | sanjana sampige | false | Medium | 2025-03-04 10:30 | 2025-03-03 |  |  | sanjana sampige |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
121 | BOM costing to be updated on the excel (VV15) | Vivid Vue |  | Completed | 100 | samiksha shetty | false | Medium | 2025-03-04 10:30 | 2025-03-05 | 2025-03-05 |  | samiksha shetty |  |  | Information received from SCM team and updated accordingly. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
122 | MAP all PRR(IDs) with test reports. | Genx12 sustenance |  | Completed | 100 | Someshwaran S;sanjana sampige | false | Medium | 2025-03-04 10:30 | 2025-02-12 | 2025-02-17 | 2025-03-04 | sanjana sampige | 1/1 | Create SVR for all sw version and Map the SPID in SVR |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
123 | unit alignment | GenX3 docs - MDD to MDR conversion |  | Completed | 100 | samiksha shetty | false | Medium | 2025-03-04 10:30 | 2025-03-05 |  |  | samiksha shetty |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
124 | Sending 3 boards to factory production | Genx3 Sustenance |  | Completed | 100 | nikhil siby | false | Medium | 2025-03-04 10:30 | 2025-03-05 |  |  | nikhil siby |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
125 | Alternate SMPS(TDK-Lambda) Check for Genx12 | Cost Productivity |  | Completed | 100 | beryl paulson | false | Medium | 2025-03-04 10:30 | 2024-07-23 01:00 | 2024-08-26 01:00 |  | Anubhav Chaudhary | 6/7 | Identify and Implementing changes to make compatible with existing M/C.;Functional Testing.;Pre-Compliance Testing at Jigani BPLMTPL Factory.;Line and Load Regulation testing.;Creating Report(Observation and Conclusion).;BOM update;Dry testing. | TDK-Lambda SMPS is received on 20th Aug 2024, the testing is in progress Dependencies:- Aman (Pre-compliance) Somesh and Nitin (BOM)(CCL) SMPS (TDK -Lambda ) is planned to release by ECN-ECR. Then only BOM update can happen. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
126 | Genx12 Field Observation(Feb 2025 validation) discussed with product team. | Genx12 sustenance |  | Completed | 100 |  | false | Medium | 2025-03-04 10:30 | 2025-03-05 | 2025-03-03 | 2025-03-04 | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
127 | Keyboard Switch Reliability | Vivid Vue |  | Completed | 100 | nikhil siby | false | Medium | 2025-03-03 10:30 | 2025-02-11 | 2025-02-10 |  | nikhil siby | 9/9 | Completed 25200 / 73000 times (11-02-2025)-(21-02-2025) 9000 times by sindhu;Completed 10600 / 73000 times (12-02-2025)(5500);Completed 52600 / 73000 times (25-02-2025)(19000) Switch tactile feedback got reduced;Completed 13100 / 73000 times (13-02-2025)(2500);Completed 33600 / 73000 times (24-02-2025)(8400);Completed 5100 / 73000 times (11-02-2025);Completed 73000 / 73000 times (03-03-2025)(9900);Completed 63100 / 73000 times (26-02-2025)(10500);Completed 16100 / 73000 times (14-02-2025)(3000) | Completed 73000 / 73000 times (03-03-2025)(9900) |  | Item | sites/ECGTeam761/Lists/Genx sustenance
128 | 26 MDD->MDR Assembly tree diagram not available for GenX3 | Genx3 Sustenance |  | Completed | 100 | Someshwaran S | false | Medium | 2025-03-03 10:30 | 2025-02-26 | 2025-02-26 |  | Someshwaran S | 1/2 | Review from Nitin Sir required;Create a assembly tree diagram -26-02-2025 |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
129 | 32 MDD->MDR GenX12- Assembly & mechanical drawings, approvals ; GenX3 - Electrical drawings, mech artworks&drwaings, assembly instructions | Genx3 Sustenance |  | Completed | 100 | Someshwaran S | false | Medium | 2025-03-03 10:30 | 2025-02-26 | 2025-03-03 | 2025-03-03 | Someshwaran S |  |  | Created assembly drawing approval note and seperate folder for assembly drawings Created mechanical drawing approval note and seperate folder for mechanical drawings Created artworks approval note and seperate folder for artworks |  | Item | sites/ECGTeam761/Lists/Genx sustenance
130 | Astute Lab machine and SMPS tracking | Genx3 Sustenance |  | Completed | 100 | Someshwaran S | false | Medium | 2025-03-03 10:30 | 2025-02-27 | 2025-02-27 | 2025-02-28 | Someshwaran S |  |  | SMPS tracking details -51617123375->Status is in transit. Machine tracking details -51608705525->Delievered on feb25 |  | Item | sites/ECGTeam761/Lists/Genx sustenance
131 | Genx3 Main PCB &Dotnet Pcb insulation discussion with Murali sir | Genx3 Sustenance |  | Completed | 100 | Anubhav Chaudhary | false | Medium | 2025-03-03 10:30 | 2025-02-28 | 2025-03-03 | 2025-03-03 | Someshwaran S |  |  | Discussed with murali sir -03-03-2025 Found mylar insulation acrylic sheet in shrikanth folder Also different shield with less thickness and height also present in shrikanth folder |  | Item | sites/ECGTeam761/Lists/Genx sustenance
132 | Display format changed- after freeze - if pressed the NEXT page/PREV page | Genx12 sustenance |  | Completed | 100 | Mahesh DE | false | Medium | 2025-03-03 10:30 | 2025-02-13 | 2025-02-14 |  | Mahesh DE | 2/3 | home screen is changed as 6x1;check the page save print freeze display;press the prev.page at menu -and take the print -print out -ok | simulation page save mode-3x4+3 - press start/stop Key- acquiring -display the FREEZE screen- press the PREV Page or NEXT page- 2/3 (3/3 actually normal condition )- press START/STOP key-after completing Print-Display the FREEZE screen -Exit -press F6 The Home screen is changed as 6x1 |  | Item | sites/ECGTeam761/Lists/Genx sustenance
133 | Updating BOM with Alternates from Kaynes | Genx12 sustenance |  | Completed | 100 | nikhil siby | false | Medium | 2025-03-02 10:30 | 2025-02-11 |  | 2025-03-03 | nikhil siby |  |  | Updated GenX12i BOM Updated GenX12i+ BOM |  | Item | sites/ECGTeam761/Lists/Genx sustenance
134 | 38 MDD->MDR Component approval note not found in GenX3 | Genx3 Sustenance |  | Completed | 100 | Someshwaran S | false | Medium | 2025-02-27 10:30 | 2025-02-26 | 2025-02-28 | 2025-02-28 | Someshwaran S |  |  | Added alternate SMPS (Longxc , emerson ) Added alternate compoinents and downloaded datasheets for the same |  | Item | sites/ECGTeam761/Lists/Genx sustenance
135 | Listing the instrument used in GenX12 document (Highlight) | Genx12 sustenance |  | Completed | 100 | beryl paulson | false | Medium | 2025-02-27 10:30 | 2025-02-10 | 2025-02-14 | 2025-02-26 | beryl paulson | 6/6 | Unit alignment;Inspection procedure;DVR;SVP ;DVP;Board alignment | 1)Did listing for -1)SVP(14/2/2025) 2)Board alignment(14/2/2025) 3)Unit alignment(14/02/2025) 2)Made report for consolidated data |  | Item | sites/ECGTeam761/Lists/Genx sustenance
136 | 30 MDD->MDR Separate usermanual approval note is not there for GenX3 | Genx3 Sustenance |  | Completed | 100 | sanjana sampige | false | Medium | 2025-02-27 10:30 | 2025-02-26 |  |  | sanjana sampige |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
137 | 31 MDD->MDR Service manual approval note not present in GenX3 | Genx3 Sustenance |  | Completed | 100 | sanjana sampige | false | Medium | 2025-02-27 10:30 | 2025-02-26 |  |  | sanjana sampige |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
138 | 46 MDD->MDR Service manual (jigani address not present) | Genx3 Sustenance |  | Completed | 100 | sanjana sampige | false | Medium | 2025-02-27 10:30 | 2025-02-26 |  |  | sanjana sampige |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
139 | Genx12(Ver1.01) testing for BTR fixed issues.(Ref:Details of Bugs in Comments Section) | Genx12 sustenance |  | Completed | 100 | samiksha shetty;sanjana sampige;Mahesh DE | false | Medium | 2025-02-26 10:30 | 2025-02-24 | 2025-02-20 | 2025-02-26 | samiksha shetty |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
140 | 28 MDD->MDR Unit alignment format is different | Genx3 Sustenance |  | Completed | 100 | samiksha shetty | false | Medium | 2025-02-26 10:30 | 2025-02-26 |  |  | samiksha shetty |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
141 | 21. MDD->MDR Cable drawings not present for Genx3 | Genx3 Sustenance |  | Completed | 100 | sanjana sampige | false | Medium | 2025-02-26 10:30 | 2025-02-26 | 2025-02-27 | 2025-02-27 | sanjana sampige |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
142 | 40 MDD->MDR Programming procedure format is different | Genx3 Sustenance |  | Completed | 100 | sanjana sampige | false | Medium | 2025-02-26 10:30 | 2025-02-26 | 2025-02-27 | 2025-02-27 | sanjana sampige |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
143 | Comparison between kitting and schematics BOM (VV15) | Vivid Vue |  | Completed | 100 | samiksha shetty | false | Medium | 2025-02-25 10:30 | 2025-02-20 |  |  | samiksha shetty |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
144 | Need to update all the cable drawings in standard format (VV15) | Vivid Vue |  | Completed | 100 | samiksha shetty | false | Medium | 2025-02-25 10:30 | 2025-01-28 |  | 2025-02-28 | samiksha shetty | 3/3 | Update the format for all the drawings;Need to be reviewed;Updating Revision history according to the previous drawings |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
145 | Test Genx12 | No Bucket |  | Completed | 100 | samiksha shetty;sanjana sampige | false | Medium | 2025-02-25 10:30 | 2025-02-20 |  |  | samiksha shetty | 2/2 | page save and LL mode;Auto and manual mode |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
146 | Update GenX3 Folder with Latest ECR379-07-2024 | Genx12 Design Transfer |  | Completed | 100 | beryl paulson | false | Medium | 2025-02-25 10:30 | 2025-02-25 | 2025-02-25 | 2025-02-25 | beryl paulson | 2/2 | DOF file -GenX3;DOF file-GenX3i | server upload |  | Item | sites/ECGTeam761/Lists/Genx sustenance
147 | BIS file task | Genx12 sustenance |  | Completed | 100 | beryl paulson | false | Medium | 2025-02-24 10:30 | 2025-02-25 | 2025-02-24 | 2025-02-24 | beryl paulson | 3/3 | RMP;RMR;User manual | TASK 1.converted each mentioned files into respective model file 2. Conversion to single pdf completed |  | Item | sites/ECGTeam761/Lists/Genx sustenance
148 | MDD->MDR = 8. Update the traceability requirements with Software verification report | Genx3 Sustenance |  | Completed | 100 | Someshwaran S;sanjana sampige | false | Medium | 2025-02-23 10:30 | 2025-02-02 | 2025-02-17 | 2025-02-21 | Anubhav Chaudhary | 0/4 | Update the tracebility requirments with Software verification report;Sample Doc Required;Update the SRS;Software requirement matrix |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
149 | 1.5 At times occasionally, even when all the leads are connected, the waveform is not displayed in any formats. GenX12i –- XQJA4G1002 | GenX12i Series observations from Camp |  | Completed | 100 | Anubhav Chaudhary;chandan kumar | false | Medium | 2025-02-23 10:30 | 2024-09-07 01:00 |  |  | Anubhav Chaudhary | 1/2 | Simulation Required (Lead off shows with lead connected);Understanding and Discussion Required | - As per our observation, it happens only in 6 channel display. - We need more clarity whether trace doesn’t appear, or straight line comes. - Not observed during our testing. Simulation Required for identifying the root cause. - Need V&V help to simulate. If get the sequence for simulation, it will great help. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
150 | 2.2 When the page marker is set to ON, the first page is detected, but printing continues blank pages until the stop button is pressed (paper wastage). GenX12i+ – YKJA4G1002 | GenX12i Series observations from Camp |  | Completed | 100 | Anubhav Chaudhary;chandan kumar | false | Medium | 2025-02-23 10:30 | 2024-09-07 01:00 |  |  | Anubhav Chaudhary | 2/3 | Simulation Required;Discussion needed;As per our observation similar issue is fixed. | - Not observed during our testing. Simulation Required for identifying the root cause. - We are trying to simulate this. If steps for simulation are provided it will be useful. Need V&V help to simulate. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
151 | 1.3 The machine during freezing in Auto mode, resulted in a blank screen with no waveforms displayed. Additionally, after freezing in the 6X2+1R display format, it reverted to the 6X1 format. GenX12i –- XQJA4G1002 | GenX12i Series observations from Camp |  | Completed | 100 | Anubhav Chaudhary | false | Medium | 2025-02-23 10:30 | 2024-09-07 01:00 |  |  | Anubhav Chaudhary | 1/2 | 6ch Issue resolved;freeze conver 6x2+1r to 6x2 instead of 6x1 | It is same as 1.1 SIMULATION REQUIRED |  | Item | sites/ECGTeam761/Lists/Genx sustenance
152 | Outcome from Discussion: Add Gap between Label and trace as it was before | GenX12i Series observations from Camp |  | Completed | 100 |  | false | Medium | 2025-02-23 10:30 | 2024-09-11 01:00 |  |  | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
153 | MDD->MDR = 2. Genx3 add Conclusion of software safety classification with intended use | Genx3 Sustenance |  | Completed | 100 |  | false | Medium | 2025-02-23 10:30 | 2025-02-02 |  |  | Anubhav Chaudhary | 1/5 | DONE ;Convert SSC to latest format. ;Software safety classification;Add conclusion to SSC with intended use;379-CE-SRC-01_SoftwareRiskClassification |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
154 | MDD->MDR = 3. add following sections in 379-Software Development Plan | Genx3 Sustenance |  | Completed | 100 | Someshwaran S | false | Medium | 2025-02-23 10:30 | 2025-02-02 | 2025-02-17 | 2025-02-18 | Anubhav Chaudhary | 1/7 | Update Soft development plan to New fromat;Task can be started;Ready for review;Add SDP ref to SDD Doc.;Create Software integration document, ;Software Verification planning.;Create Software integration test plan Doc. |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
155 | MDD->MDR = 4. add following sections in 379-SRS-01_Rev_01 Software Requirement Specification.pdf | Genx3 Sustenance |  | Completed | 100 | Someshwaran S | false | Medium | 2025-02-23 10:30 | 2025-02-02 | 2025-02-10 | 2025-02-14 | Anubhav Chaudhary | 3/5 | Requirement requirement analysis;Risk Control Measure.;Communication requirements.;Software System inputs and outputs.;Software Requirement (Development tool, programming language, debugging tool). |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
156 | MDD->MDR = 7. share updated Software Problem Resolution Process document QARA | Genx3 Sustenance |  | Completed | 100 | Someshwaran S | false | Medium | 2025-02-23 10:30 | 2025-02-02 | 2025-02-17 | 2025-02-21 | Anubhav Chaudhary | 3/4 | Share the Updated PRR Doc.;Update the PRR Doc to new format;Update the content to PRR Doc;ready for review |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
157 | MDD->MDR = 10. Software Configuration Management Process document | Genx3 Sustenance |  | Completed | 100 | Someshwaran S | false | Medium | 2025-02-23 10:30 | 2025-02-03 | 2025-02-21 | 2025-02-21 | Anubhav Chaudhary | 0/2 | Clause 8: Software configuration management process;Sample Doc |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
158 | MDD-MDR = 11. updated software release notes with complete modification changes/release version with Reference: 379-UVR-01_GENX3 Unit Test results2 (Unit Test of Genx3 series) | Genx3 Sustenance |  | Completed | 100 |  | false | Medium | 2025-02-23 10:30 | 2025-02-03 |  |  | Anubhav Chaudhary | 0/1 | Clause 5.8: Software version control release note |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
159 | MDD->MDR = 12. PrepareSystem Integration plan and System integration test report | Genx3 Sustenance |  | Completed | 100 | Someshwaran S;sanjana sampige | false | Medium | 2025-02-23 10:30 | 2025-02-03 | 2025-02-21 | 2025-02-21 | Anubhav Chaudhary | 4/6 | Create System Integration plan and System integration test report;As per Clause 5.8: Software integration plan and report ;Sample Doc Req;Clause 5.8: Software integration plan and report ;Software integration plan;Review |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
160 | MDD->MDR = 13. update Software Validation & Verification Report with respect to the following Plan 379_VEP-05_DesignVerificationPlan +L75 | Genx3 Sustenance |  | Completed | 100 | chandan kumar | false | Medium | 2025-02-23 10:30 | 2025-02-03 |  |  | Anubhav Chaudhary | 0/2 | Clause 5.1.6: Software V&V plan;update Software Validation & Verification Report with respect to 379_VEP-05_DesignVerificationPlan |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
161 | MDD->MDR = 14. update Software Validation & Verification Report with respect to the following reports 379_VER-05_Rev_B_DesignVerificationReport | Genx3 Sustenance |  | Completed | 100 | chandan kumar | false | Medium | 2025-02-23 10:30 | 2025-02-03 |  |  | Anubhav Chaudhary | 0/2 | update Software Validation & Verification Report with respect to 379_VEP-05_DesignVerificationRepor;Software Verification report |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
162 | MDD->MDR = 15.Clause 6 : Maintanance process of the product software to be Justified in the 62304 test Report by BPL | Genx3 Sustenance |  | Completed | 100 | sanjana sampige | false | Medium | 2025-02-23 10:30 | 2025-02-03 | 2025-02-20 | 2025-02-20 | Anubhav Chaudhary | 0/1 | Software maintenance process |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
163 | MDD->MDR = 16. updated as per EN IEC 62304:2006+AMD1:2015 by BPL | Genx3 Sustenance |  | Completed | 100 | Someshwaran S | false | Medium | 2025-02-23 10:30 | 2025-02-03 | 2025-02-20 | 2025-02-20 | Anubhav Chaudhary | 1/3 | READY FOR REVIEW ;Task Created file available https://bplmedtech-my.sharepoint.com/:f:/g/personal/anubhav_chaudhary_b ;EN 62304 test report |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
164 | Sending Genx3 mahine to Astute with Longxc Smps | Genx3 Sustenance |  | Completed | 100 | Someshwaran S;Nitin Chavan | false | Important | 2025-02-20 10:30 | 2025-02-18 |  | 2025-02-19 | Someshwaran S |  |  | Assemble one Genx3 Mchine with longxc SMPS Send the same to Astute lab for compliance testing |  | Item | sites/ECGTeam761/Lists/Genx sustenance
165 | Inspection Procedure | GenX3 docs - MDD to MDR conversion |  | Completed | 100 | samiksha shetty | false | Medium | 2025-02-19 10:30 | 2025-02-20 |  | 2025-02-20 | samiksha shetty |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
166 | Print after Paper feed key press with black mark on | Genx12 sustenance |  | Completed | 100 | Mahesh DE | false | Medium | 2025-02-19 10:30 | 2025-01-31 | 2025-02-19 | 2025-02-20 | Mahesh DE | 2/2 | check the related Function();check the simulation steps |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
167 | Display format changed- after freeze - if Print format changed | Genx12 sustenance |  | Completed | 100 | Mahesh DE | false | Medium | 2025-02-18 10:30 | 2025-02-13 | 2025-02-13 | 2025-02-18 | Mahesh DE | 2/2 | check the home screen- take the print out-on FREEZE screen-There is no problem-at normal condition;if changed the PRINT FORMAT - then the home screen is changes to 6x1 | The home screen is changed - after freeze steps- 1- home screen is 3x4+2 - page save- mode- press start/stop key- after acquiring- Freeze screen appeared on display screen- step 2- change the Print format- and press the START/STOP key exit the freeze screen - main screen is 6x1 |  | Item | sites/ECGTeam761/Lists/Genx sustenance
168 | freeze Page number is Mis matched | Genx12 sustenance |  | Completed | 100 | Mahesh DE | false | Medium | 2025-02-18 10:30 | 2025-02-13 | 2025-02-12 | 2025-02-12 | Mahesh DE | 6/6 | taken log of after filtered data and check the sampling rate;check the Freeze Page Function- add the few condition;check the freeze print and display pages ;check the all sampling data -;check the other MODE;taken log of the before filtered data | press -freeze Screen at 3x4+2 home screen -Pages showing as 2/3 ,but it should be as 3/3 |  | Item | sites/ECGTeam761/Lists/Genx sustenance
169 | Coordinating with Kaynes For RT view issue | Genx12 sustenance |  | Completed | 100 | Anubhav Chaudhary;Someshwaran S;nikhil siby | false | Medium | 2025-02-17 10:30 | 2025-02-13 |  | 2025-02-18 | Anubhav Chaudhary |  |  | Got the 2 bridge PCB and 2 com.int pcb form kaynes on 15-2-25 Checked the RT View,FTP, USB transfer and its working properly-15-2-2025 Put the 1 machine in Auto FTP upload-17-2-25 Put the 1 machine in PC Mode (RT-View)-17-2-25 Boards functioned properly and need to return the baords -18-02-2025 |  | Item | sites/ECGTeam761/Lists/Genx sustenance
170 | GENX3 CE Bom queries | Genx3 Sustenance |  | Completed | 100 | Someshwaran S;Nitin Chavan | false | Important | 2025-02-16 10:30 | 2025-02-17 | 2025-02-15 | 2025-02-17 | Someshwaran S |  |  | For genx3-CE Main Board some queries came for EMS vendor need to clear them |  | Item | sites/ECGTeam761/Lists/Genx sustenance
171 | Review and update Genx3 documents-SAD,SMP,SDD | Genx3 Sustenance |  | Completed | 100 | Anubhav Chaudhary;Someshwaran S;sanjana sampige | false | Medium | 2025-02-13 10:30 | 2025-02-14 | 2025-02-14 | 2025-02-14 | Someshwaran S |  |  | Review and updated the SAD,SMP,SDD |  | Item | sites/ECGTeam761/Lists/Genx sustenance
172 | Arrange/separate 15 cables for the next 15 trial units.(VV15) | Vivid Vue |  | Completed | 100 | samiksha shetty | false | Medium | 2025-02-13 10:30 | 2025-01-28 |  |  | samiksha shetty |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
173 | Release Genx3 ECN | To do |  | Completed | 100 | nikhil siby | false | Important | 2025-02-13 10:30 | 2024-08-23 01:00 | 2024-11-01 01:00 | 2025-02-11 | nikhil siby | 7/7 | Testing GenX3 obsolete components;V&V request for testing;Updating BOM ;Request mahesh for price impact report;Design review meeting;Share documents with QARA;Creating test report | Discussed with mahesh he will initiate ECR GenX3 BOM is having some issue related to SAP code mismatch. After somesh clearing this issue with the help of faris BOM can be updated. Updated BOM but making corrections as per the QARA comments Requested to faris for new SAP codes Requested mahesh for price impact report and stock details. After receiving the details ECN can be closed.(11-02-25) Price impact report is received from mahesh but stock details didnt received(12-02-25) stock details received from dhanush(12-02-25) shared all documents with QARA(12-02-25) ECN will be released by 13-02-2025 |  | Item | sites/ECGTeam761/Lists/Genx sustenance
174 | Genx12 Fact Dot Net Board Debug(Voltage) | Genx12 sustenance |  | Completed | 100 | Someshwaran S;sanjana sampige | false | Medium | 2025-02-13 10:30 | 2025-01-28 | 2025-02-10 | 2025-02-12 | Someshwaran S | 7/7 | 5 Board with Factory .;Check the test point voltages.;Reprogram the board and send to factory;3 boards with Store.;Check shield shorting with the 5 PIN Debug connector.;Getting new connector;getting new shield | As shield is getting short ,so removed the shield in One board and tested ,board getting booted and able to load the sw. So bend the part of the shield where it is getting short and checked ,board getting booted . Replaced the current shield and current Programming connector with new one and updated the boards with latest software. Completed the basic level of testing(USB,FTP) |  | Item | sites/ECGTeam761/Lists/Genx sustenance
175 | Add RT view to Software Documents | No Bucket |  | Completed | 100 | sanjana sampige | false | Medium | 2025-02-13 10:30 | 2025-02-14 |  | 2025-02-14 | sanjana sampige |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
176 | 6 NOS FACTORY Dotnet board not booting-debug &clear | Genx12 sustenance |  | Completed | 100 | Someshwaran S;sanjana sampige | false | Medium | 2025-02-13 10:30 | 2025-01-30 | 2025-02-10 | 2025-02-12 | Someshwaran S | 0/3 | Get new connector and Shiled ;Send back to HO;Fix them and Reprogram the board | As shield is getting short ,so removed the shield in One board and tested ,board getting booted and able to load the sw. So bend the part of the shied where it is getting short and checked ,board getting booted .Replaced the new programming and Shield with on existing one. Flashed the latest software and Verified the basic functions(FTP.USB) |  | Item | sites/ECGTeam761/Lists/Genx sustenance
177 | DTL GenX3 | Genx3 Sustenance |  | Completed | 100 | nikhil siby | false | Medium | 2025-02-12 10:30 | 2025-01-20 |  |  | nikhil siby |  |  | DTL Updated copy |  | Item | sites/ECGTeam761/Lists/Genx sustenance
178 | Power PCB Debug Charging indication after burning. | Genx12 sustenance |  | Completed | 100 | nikhil siby | false | Medium | 2025-02-12 10:30 | 2025-01-28 |  | 2025-02-14 | nikhil siby | 3/3 | Try Change the output capacitor.;Complete discharge and charging test(1 iteration);Change the high drive and drive(Take CWS help) | After changing all the components the board is working Kept for complete discharge and charging test Sent boards to factory and handed over to store |  | Item | sites/ECGTeam761/Lists/Genx sustenance
179 | Attaching ECR of GenX3 Document to Design transfer list | Genx12 sustenance |  | Completed | 100 | beryl paulson | false | Medium | 2025-02-12 10:30 | 2025-02-10 | 2025-02-13 | 2025-02-13 | beryl paulson | 4/4 | Attachment of GenX3 ECR sl no label;Attachment of GenX3 ECR sl no label; attachment of GenX3ECR long lead document to main folder;attachment of GenX3ECR long lead document to main folder | 1)Attaching ECR Long lead document to main folder 13/2/2025 done 2)Completed on 13/2/2025 |  | Item | sites/ECGTeam761/Lists/Genx sustenance
180 | In freeze after printing spike observed on display. Sometimes | Genx12 sustenance |  | Completed | 100 | Mahesh DE | false | Medium | 2025-02-12 10:30 | 2025-01-31 | 2025-02-10 | 2025-02-13 | Mahesh DE | 6/6 | implement- add the few function and IF -condition;check the page save mode and take the print out -and the there is no spike on the FREEZE screen ;Analyzing and Identification - the spike at freeze - most of the time 3x4+2 and 6x2+1 ;Take the log without filter- There is no spike;Simulation;plan fix- check the Freeze function , Filter_. and Related function - taken LOG with filter | the freeze-spike it is mainly appeared at the 3x4+2and 6x2+1 ,and at the time page number is showing as 2/3- next page -F4 key not working |  | Item | sites/ECGTeam761/Lists/Genx sustenance
181 | Vivid vue battery charging time analysis | Vivid Vue |  | Completed | 100 | beryl paulson | false | Medium | 2025-02-12 10:30 | 2025-01-31 | 2025-01-29 | 2025-02-14 | beryl paulson | 4/4 | Discharging battery 1;Charging battery in progress;Completed rework and finished charging in battery (4);Did rework for charging current 1A | Completed for 1A charging current |  | Item | sites/ECGTeam761/Lists/Genx sustenance
182 | Genx12 factory 3 Boards RT- View Communication Issue | Genx12 sustenance |  | Completed | 100 | Anubhav Chaudhary;Someshwaran S | false | Medium | 2025-02-11 10:30 | 2025-02-12 |  | 2025-02-11 | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
183 | Document verification for Internal Audit GenX12 | No Bucket |  | Completed | 100 | beryl paulson | false | Important | 2025-02-10 10:30 | 2025-02-07 | 2025-02-07 |  | beryl paulson | 10/10 | Usability Verification (470);DRR ;Software anomalies report;Reviewd GenX12i ECN folder(7/2/2025);Usability validation(469,470);Review of GenX12 files;Feasibility ;NC folder creation and file maintenance;Hardware design document ;Test reports | 1)Started NC folder creation 2)started review of documents Review in progress..... |  | Item | sites/ECGTeam761/Lists/Genx sustenance
184 | Passive alternate components for kaynes | Genx12 sustenance |  | Completed | 100 | nikhil siby | false | Medium | 2025-02-10 10:30 | 2025-02-11 |  | 2025-02-11 | nikhil siby |  |  | 1. 2439BEK297 Make: Vishay MPN: CRCW06033K90FKEA alternate: RK73H1JTTD3901F Make: KOA 1. 2439BEK381 actual make: Walsin : 0402B104K250CT alternate: ​CGA2B3X7R1E104K050BB TDK |  | Item | sites/ECGTeam761/Lists/Genx sustenance
185 | Document review of DPF | Genx12 sustenance |  | Completed | 100 | Sindhu Kashyap;beryl paulson | false | Medium | 2025-02-09 10:30 | 2025-02-10 |  |  | beryl paulson | 17/17 | 3.DESIGN INPUT REQUIREMENT;15.DESIGN FEASIBILITY REPORT;13.SOFTWARE VERIFICATION PLAN;7.PEMS ARCHITECTURE (Programmable Electrical Medical Systems);8.PROBLEM RESOLUTION REPORT;11.SOFTWARE VALIDATION REPORT;6.CODE REVIEW REPORT;14.SOFTWARE ANOMALIES REPORT;5.SOFTWARE SAFETY CLASSIFICATION;10.SOFTWARE DEVELOPMENT PLAN;17.HARDWARE DESIGN DOCUMENT;9.SOFTWARE ARCHITECTURE;1.PNPD(PRODUCT NEED AND PROPOSAL DOCUMENT);12.SOFTWARE VERIFICATION REPORT;2.DESIGN PROJECT PLAN;16.DRR;4.DESIGN USABILITY REPORT |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
186 | Attaching ECR-469-01-2025 in Latest Design transfer document file | Genx12 sustenance |  | Completed | 100 | beryl paulson | false | Important | 2025-02-09 10:30 | 2025-02-10 | 2025-02-10 | 2025-02-11 | beryl paulson | 1/1 | Add ECR document in folder (QNAP) |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
187 | GenX3 All Docs old to new format gap assessment . | Genx3 Sustenance |  | Completed | 100 | sanjana sampige | false | Medium | 2025-02-09 10:30 | 2025-02-05 |  |  | sanjana sampige | 2/2 | Create excel doc with gap information;Identify which documents have gaps |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
188 | GenX12 production requirement 06/02/2025 | Genx12 sustenance |  | Completed | 100 | beryl paulson | false | Urgent | 2025-02-06 10:30 | 2025-02-07 |  |  | beryl paulson | 2/2 | Did basic test 6/2/2025;Board assembling in production |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
189 | Ld Placement Image update User manual and transfer QNS Activity | Genx12 sustenance |  | Completed | 100 | beryl paulson | false | Medium | 2025-02-05 10:30 | 2025-02-06 |  |  | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
190 | Lead Placement Image update->Code | Genx12 sustenance |  | Completed | 100 | Anubhav Chaudhary | false | Medium | 2025-02-05 10:30 | 2025-02-06 |  |  | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
191 | transfer 500 records file few records not able to transfer | To do |  | Completed | 100 | Mahesh DE | false | Medium | 2025-02-05 10:30 | 2024-07-18 01:00 | 2025-01-23 | 2025-01-30 | Mahesh DE | 1/1 | FTP transfer Only |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
192 | visibility study | New Platform Tasks |  | Completed | 100 |  | false | Medium | 2025-02-05 10:30 | 2025-02-06 |  |  | Mahesh DE |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
193 | Rework for 3 Tech boards(VV15) | Vivid Vue |  | Completed | 100 |  | false | Medium | 2025-02-05 10:30 | 2025-02-06 |  | 2025-02-04 | samiksha shetty |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
194 | Dot Net Board Programming procedure update | Genx12 Design Transfer |  | Completed | 100 | Someshwaran S | false | Medium | 2025-02-05 10:30 | 2025-01-28 |  |  | Someshwaran S |  |  | Updated Dotnet board programming procedure with terminal image and added SW version check in procedure |  | Item | sites/ECGTeam761/Lists/Genx sustenance
195 | Mount the removed components on 2 keypad boards and flash them.(VV15) | Vivid Vue |  | Completed | 100 | samiksha shetty | false | Medium | 2025-02-04 10:30 | 2025-01-28 |  |  | samiksha shetty |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
196 | Genx3 Printer Check Alternate | Cost Productivity |  | Completed | 100 | Someshwaran S | false | Medium | 2025-02-04 10:30 | 2024-07-23 01:00 | 2024-08-05 01:00 | 2024-09-11 01:00 | Anubhav Chaudhary | 5/7 | Perform Speed Test.;Awaiting for Samples.;Perform printer test.;Check with print dark and normal setting;Create test report ;Take print with maximum Gain and higher heart rate.(For Maximum Current);Compare print quality between Fujitsu and PRT (Attach Proof with report) | Put ON HOLD (Price High) Dependency-Inder Sir and SCM Inder Sir ask to hold as he is working with vendor to get new samples and Cost productivity. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
197 | Genx3 LCD Alternate | Cost Productivity |  | Completed | 100 | Someshwaran S | false | Medium | 2025-02-04 10:30 | 2024-07-23 01:00 |  |  | Anubhav Chaudhary | 1/3 | Compare the alternate lcd specs with exisiting lcd;Getting Samples.;Testing. | PRICE is HIGH so put on HOLD Dependency-SCM and Inder Sir. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
198 | Genx3 CE Bom in SAP | Genx3 Sustenance |  | Completed | 100 | Someshwaran S | false | Medium | 2025-02-04 10:30 | 2025-02-05 |  | 2025-02-05 | Someshwaran S |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
199 | DPL Review in template | Vivid Vue |  | Completed | 100 | Someshwaran S | false | Medium | 2025-02-03 10:30 | 2025-01-28 |  | 2025-01-28 | Someshwaran S |  |  | Completed for all the PCBA |  | Item | sites/ECGTeam761/Lists/Genx sustenance
200 | Create MSP430 Programming Procedure Document for Genx12 | Genx12 sustenance |  | Completed | 100 | Someshwaran S | false | Medium | 2025-02-03 10:30 | 2025-01-29 |  |  | Someshwaran S |  |  | Created MSP430 programming procedure. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
201 | Updating ECN DOF, DTL based on review from QARA | Genx12 sustenance |  | Completed | 100 | beryl paulson | false | Important | 2025-02-02 10:30 | 2025-01-31 | 2025-01-31 |  | beryl paulson |  |  | QARA shared feedback Updating DOF,DTL ECN completed as per QARA remarks |  | Item | sites/ECGTeam761/Lists/Genx sustenance
202 | User manual review with print taken | Genx12 sustenance |  | Completed | 100 | Sindhu Kashyap;beryl paulson | false | Important | 2025-02-02 10:30 | 2025-01-31 |  |  | beryl paulson | 4/4 | Review done from PC;other print Beryl cross checked to confirm no mistake from R&D;One print Sindhu is crosss checking;Taken print out physically 3 times | Review done from PC Taken print out physically One print Sindhu is crosss checking other print Beryl cross checked to confirm no mistake from R&D This print will be our reference for future. Send to SCM final copy |  | Item | sites/ECGTeam761/Lists/Genx sustenance
203 | Keyboard Version sometimes in some factory boards observed as .170 | Genx12 sustenance |  | Completed | 100 | Anubhav Chaudhary | false | Medium | 2025-02-02 10:30 | 2025-01-31 |  | 2025-02-03 | Anubhav Chaudhary | 5/5 | Read command.;Touch was done on connector and Passive components on TX and RX line.;Factory 5 Boards this was observed;With code modification this was reduced to rare .;Receive_Keyboard_version_lock_flag was created and Read key() function was blocked during Version |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
204 | Ethernet Link Down (Machine did not restart - by chandan) | Genx12 sustenance |  | Completed | 100 | Anubhav Chaudhary | false | Medium | 2025-02-02 10:30 | 2025-01-31 |  |  | Anubhav Chaudhary | 0/1 | Tested the same unit on 01/02/2025 .All(800) Records transferred. |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
205 | ECN impact analysis stock details entry | Genx12 sustenance |  | Completed | 100 | beryl paulson | false | Important | 2025-01-30 10:30 | 2025-01-31 | 2025-01-30 | 2025-01-31 | beryl paulson |  |  | As per QARA input Stock should be maintained in ECN based on confirmation from stock. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
206 | Genx12 user manual (check) | Genx12 Design Transfer |  | Completed | 100 | beryl paulson | false | Medium | 2025-01-30 10:30 | 2025-01-28 |  |  | beryl paulson | 0/3 | Check fig & reference.;Check table reference.;Font Check (Aller light is standard) |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
207 | Store Material Record alignment. | No Bucket |  | Completed | 100 | Anubhav Chaudhary;Someshwaran S;beryl paulson;nikhil siby;Mahesh DE | false | Medium | 2025-01-30 10:30 | 2025-01-27 | 2025-01-27 | 2025-01-28 | Someshwaran S |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
208 | Reflashing Dot Net Boards 47 numbers at factory | Genx12 sustenance |  | Completed | 100 | Anubhav Chaudhary;Someshwaran S | false | Medium | 2025-01-30 10:30 | 2025-01-31 |  | 2025-01-23 | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
209 | Reflash programming in Main,Key,ECG | Genx12 sustenance |  | Completed | 100 | Anubhav Chaudhary;Someshwaran S | false | Medium | 2025-01-30 10:30 | 2025-01-30 |  | 2025-01-30 | Anubhav Chaudhary |  |  | Main Board-24no's ECG-25NO'S Key-25no's |  | Item | sites/ECGTeam761/Lists/Genx sustenance
210 | Working on QARA feedback for DIR and brochure | Genx12 sustenance |  | Completed | 100 | nikhil siby | false | Medium | 2025-01-30 10:30 | 2025-01-20 |  |  | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
211 | Create ECR Design Review Report | Genx12 Design Transfer |  | Completed | 100 | beryl paulson | false | Medium | 2025-01-29 10:30 | 2025-01-28 |  |  | beryl paulson |  |  | 1. Same code(379) is used across Genx3,Genx3i and Genx12i Series. 2. For Genx12i existing deck sheets in stock will be consumed first. 3. For Genx12i+ new trade mark will be implemented. 4. Jan 2025 Assembled Genx12i will be used with same deck sheet. 5. Genx12i+ Painted cabinets will not be used. 6. Stock: 7. 1. cushion 50 8. 2. Deck Sheet 213 9. 3. User Manual 50 10. 4. Carton Box 50 11. 5. Paper Support 150 12. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
212 | while trnsfer 500 record to FTP- middle of the Transfer ,unit is hang | Genx3 Sustenance |  | Completed | 100 | Mahesh DE | false | Medium | 2025-01-28 10:30 | 2025-01-16 | 2025-01-16 | 2025-01-23 | Mahesh DE | 4/4 | Debug simulation step compare to the Genx12 ;After Delay Show the REVIVE message;add the DOT NET Restart Function;check ping connection | Simulation steps: 1.Check the PING - connected 2.option- mark all record 3. while thransferring all record- press the BACK key 1. unit will showing as TWO Icon 2. it will not come back to Home screen |  | Item | sites/ECGTeam761/Lists/Genx sustenance
213 | while trnffering record to FTP- some records are not transferred | Genx3 Sustenance |  | Completed | 100 | Mahesh DE | false | Medium | 2025-01-28 10:30 | 2025-01-16 | 2025-01-13 | 2025-01-22 | Mahesh DE | 5/5 | check the PC not responding/ server not responding;connect the debugger and check the debug point;store 500 record;5 times transfer for FTP;check the Record is it LL mode data | Transfer all record - 1 struck/hang the unit 1. some records are not transferred |  | Item | sites/ECGTeam761/Lists/Genx sustenance
214 | while transfering Data to FTP- The Unit is hang | Genx3 Sustenance |  | Completed | 100 | Mahesh DE | false | Medium | 2025-01-28 10:30 | 2025-01-16 | 2025-01-16 | 2025-01-20 | Mahesh DE | 8/8 | Ping the CMD and check the 192.168.10.11 server ;if Record not transfer- at the time DOT net board Restart.;connect the UART cable to DOT NET board and check the Initialization at Tera term ;store the 500 records with interpreatation ;Upload Records and send through FTP;Mark all Record and Send through FTP;" Dotnet Reboot" function is added -and after transfer100 records automatically Dotnet Board Restart;check the FTP connectivity |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
215 | Processing Msg update(Chandan's BTR) | Genx12 sustenance |  | Completed | 100 | Mahesh DE | false | Medium | 2025-01-28 10:30 | 2025-01-16 | 2025-01-27 | 2025-01-29 | Mahesh DE | 5/5 | check all the testing MODE -and print out - working fine;at Page save check the parameters -speed and added one condition at Print state -ECG_PRINT_IDLE ;first take the -Recent version and simulate the process along with chandan;added One condition for -"Print state" and added clear screen function-Processing message complete;initial stage -check the speed | Simulation:::Pacer ON-> Freeze->Print->F1->Print finishes->Change Speed to 50->Print(F1) |  | Item | sites/ECGTeam761/Lists/Genx sustenance
216 | Kitting BOM datsheet finding | Vivid Vue |  | Completed | 100 | samiksha shetty;beryl paulson;nikhil siby | false | Medium | 2025-01-28 10:30 | 2025-01-20 | 2025-01-14 |  | samiksha shetty |  |  | ITC, MIP ,Alarm PCB completed |  | Item | sites/ECGTeam761/Lists/Genx sustenance
217 | Update latest MTER in DT folder | Genx12 Design Transfer |  | Completed | 100 | beryl paulson | false | Medium | 2025-01-28 10:30 | 2025-01-28 |  | 2025-01-28 | beryl paulson |  |  | updated in folder |  | Item | sites/ECGTeam761/Lists/Genx sustenance
218 | Carton image & BOM update | Genx12 Design Transfer |  | Completed | 100 | Someshwaran S;beryl paulson;Ankar Das | false | Medium | 2025-01-28 10:30 | 2025-01-28 |  | 2025-01-28 | Ankar Das | 0/4 | Image update by-Ankar;ECN Impact Update -Beryl;BOM update-Somesh;SAP Code generation-Somesh |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
219 | Test Genx3 Machine connectivity functionality with Dot Net board(Ver 1.9) and create report) | Genx12 Design Transfer |  | Completed | 100 | Sindhu Kashyap | false | Medium | 2025-01-27 10:30 | 2025-01-27 |  |  | Sindhu Kashyap | 9/9 | 7. Upload ECG on FTP server.;5. Export the Recorded ECG.;9. Auto FTP Upload.;2. USB Flash Drive test.(Check the ECG morphology, data and S/W version );3. USB Printer test.;4. Ping test;1. Self test ALL;8. Auto USB Upload.;6. Print ECG on USB Printer. |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
220 | 50 Hz Notch Filter Plan and design. | Feature Addition |  | Completed | 100 | samiksha shetty | false | Medium | 2025-01-27 10:30 | 2024-10-30 01:00 |  | 2025-01-31 | samiksha shetty |  |  | HPF Freq(cut) == 51 Hz LPF Freq(cut) == 49 Hz |  | Item | sites/ECGTeam761/Lists/Genx sustenance
221 | Battery connector comparison (VV15) | Vivid Vue |  | Completed | 100 | samiksha shetty | false | Medium | 2025-01-27 10:30 | 2025-01-28 |  |  | samiksha shetty |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
222 | 8. Trademark Change - Brochure | Genx12 Design Transfer |  | Completed | 100 | samiksha shetty | false | Medium | 2025-01-27 10:30 | 2025-01-22 |  | 2025-01-24 | samiksha shetty | 1/1 | Include 3 changes in Brochure as per feedback from QARA.(Ref: Sindhu's report) | As per Marketing team-by EOD(27/01/2025) it will given. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
223 | Prepare ECN for Genx12 | Genx12 Design Transfer |  | Completed | 100 | beryl paulson | false | Medium | 2025-01-27 10:30 | 2025-01-24 |  |  | beryl paulson | 1/2 | ECG Impact Analysis;ECN(90%) | ECN prepared and reviwed |  | Item | sites/ECGTeam761/Lists/Genx sustenance
224 | 9. Trademark Change - DOF, DTL | Genx12 Design Transfer |  | Completed | 100 | Sindhu Kashyap;beryl paulson | false | Medium | 2025-01-26 10:30 | 2025-01-22 |  | 2025-01-24 | Sindhu Kashyap | 0/3 | Update Doc number of Assembly Label approval Note(469-ALA-001);Update Doc number of Assembly Label approval Note(470-ALA-001);Beryl's Review (Comparison) |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
225 | 18. Genx12 Label instruction Document | Genx12 Design Transfer |  | Completed | 100 | nikhil siby | false | Medium | 2025-01-26 10:30 | 2025-01-22 |  | 2025-01-23 | nikhil siby | 2/2 | Received from Ankar.. Our Review Done;Girish Review Pending.(We included it in Design Transfer Folder) |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
226 | Mech Approval Note updation | Genx12 Design Transfer |  | Completed | 100 | Someshwaran S | false | Medium | 2025-01-26 10:30 | 2025-01-24 |  | 2025-01-27 | Anubhav Chaudhary | 0/1 | Completed by Somesh (Need to include in folder) | Completed by Somesh (Need to include in folder) |  | Item | sites/ECGTeam761/Lists/Genx sustenance
227 | Assembly Approval Note updation | Genx12 Design Transfer |  | Completed | 100 | nikhil siby | false | Medium | 2025-01-26 10:30 | 2025-01-24 |  | 2025-01-27 | Anubhav Chaudhary |  |  | Completed By Nikhil(Need to push in Repository) |  | Item | sites/ECGTeam761/Lists/Genx sustenance
228 | 6. Trademark Change - User Manual | Genx12 Design Transfer |  | Completed | 100 | beryl paulson | false | Medium | 2025-01-26 10:30 | 2025-01-22 | 2025-01-22 | 2025-01-24 | Anubhav Chaudhary |  |  | Under Nitin Sir's Review |  | Item | sites/ECGTeam761/Lists/Genx sustenance
229 | 10. Trademark Change - Label Drawing Approval Notes | Genx12 Design Transfer |  | Completed | 100 | Someshwaran S | false | Medium | 2025-01-26 10:30 | 2025-01-22 |  | 2025-01-24 | Someshwaran S | 0/1 | Completed |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
230 | 11. Trademark Change - SAP code update | Genx12 Design Transfer |  | Completed | 100 | Someshwaran S | false | Medium | 2025-01-26 10:30 | 2025-01-22 |  | 2025-01-24 | Someshwaran S | 2/2 | 1. Mechanical Dependency (Jaswin);2. Share with Farris |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
231 | 12. Trademark Change - DPL | Genx12 Design Transfer |  | Completed | 100 | Someshwaran S | false | Medium | 2025-01-26 10:30 | 2025-01-22 |  | 2025-01-24 | Someshwaran S | 1/1 | Mechanical Dependency (Jaswin) |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
232 | 13. Trademark Change - SAP BOM | Genx12 Design Transfer |  | Completed | 100 | Someshwaran S | false | Medium | 2025-01-26 10:30 | 2025-01-22 |  | 2025-01-24 | Someshwaran S | 1/1 | Mechanical Dependency (Jaswin) | Mechanical Dependency (Jaswin) |  | Item | sites/ECGTeam761/Lists/Genx sustenance
233 | 4. Trademark Change - Display Assembly | Genx12 Design Transfer |  | Completed | 100 | Someshwaran S | false | Medium | 2025-01-23 10:30 | 2025-01-22 |  | 2025-01-24 | Anubhav Chaudhary | 0/1 | Mechanical Dependency (Jaswin) | As Mech team(Jaswin) does not capture the revision. So this task is not required. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
234 | 5. Trademark Change - Unit Assembly | Genx12 Design Transfer |  | Completed | 100 | nikhil siby | false | Medium | 2025-01-23 10:30 | 2025-01-22 |  | 2025-01-24 | Anubhav Chaudhary | 0/1 | Mechanical Dependency (Jaswin) | As Mech team(Jaswin) does not capture the revision. So this task is not required. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
235 | 14. Trademark Change - Cushion | Genx12 Design Transfer |  | Completed | 100 | Someshwaran S | false | Medium | 2025-01-23 10:30 | 2025-01-22 |  | 2025-01-24 | Anubhav Chaudhary | 0/1 | Mechanical Dependency (Jaswin) |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
236 | 3. Trademark Change - Assembly Drawing | Genx12 Design Transfer |  | Completed | 100 | nikhil siby | false | Medium | 2025-01-23 10:30 | 2025-01-22 |  | 2025-01-24 | Anubhav Chaudhary | 0/1 | Mechanical Dependency (Jaswin) | As Mech team(Jaswin) does not capture the revision. So this task is not required. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
237 | GenX12 User manual updation | Genx12 Design Transfer |  | Completed | 100 | beryl paulson | false | Medium | 2025-01-23 10:30 | 2025-01-20 |  | 2025-01-20 | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
238 | 17. Genx3 SRN for Dot Net Board | Genx12 Design Transfer |  | Completed | 100 | Sindhu Kashyap;Anubhav Chaudhary | false | Medium | 2025-01-23 10:30 | 2025-01-22 |  | 2025-01-24 | Anubhav Chaudhary |  |  | Completed for Genx3 and Genx3i |  | Item | sites/ECGTeam761/Lists/Genx sustenance
239 | 19. User manual Approval Note | Genx12 Design Transfer |  | Completed | 100 | beryl paulson | false | Medium | 2025-01-23 10:30 | 2025-01-24 |  |  | Anubhav Chaudhary | 0/1 | Completed |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
240 | 20. Service manual Approval note | Genx12 Design Transfer |  | Completed | 100 | beryl paulson | false | Medium | 2025-01-23 10:30 | 2025-01-24 |  |  | Anubhav Chaudhary | 0/1 | Completed |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
241 | 7. Trademark Change - Service Manual | Genx12 Design Transfer |  | Completed | 100 | Sindhu Kashyap;nikhil siby | false | Medium | 2025-01-23 10:30 | 2025-01-22 |  | 2025-01-24 | Anubhav Chaudhary |  |  | Shared with Beryl |  | Item | sites/ECGTeam761/Lists/Genx sustenance
242 | 16. Trademark Change - IPD | Genx12 Design Transfer |  | Completed | 100 | samiksha shetty;nikhil siby | false | Medium | 2025-01-23 10:30 | 2025-01-22 |  | 2025-01-24 | Anubhav Chaudhary |  |  | Shared with Beryl(Completed) |  | Item | sites/ECGTeam761/Lists/Genx sustenance
243 | Add Paper feed with and without Black mark ON in SRS and SVP | Genx12 sustenance |  | Completed | 100 | Sindhu Kashyap | false | Medium | 2025-01-22 10:30 | 2025-01-16 |  |  | Sindhu Kashyap |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
244 | Freeze Print all data with and without Black Mark detection. | Genx12 sustenance |  | Completed | 100 | Anubhav Chaudhary | false | Medium | 2025-01-22 10:30 | 2025-01-16 |  | 2025-01-23 | Anubhav Chaudhary | 4/4 | PRINTING PRINT ALL FREEZE DATA (Without Black Mark Detection);PRINTING PRINT ALL FREEZE DATA( With Black Mark ON);DISPLAY PRINT ALL FREEZE DATA (Without Black Mark Detection);DISPLAY PRINT ALL FREEZE DATA (With Black Mark Detection) |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
245 | 15 Trademark Change - UAT | Genx12 Design Transfer |  | Completed | 100 | samiksha shetty | false | Medium | 2025-01-21 10:30 | 2025-01-22 |  | 2025-01-24 | samiksha shetty |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
246 | 1. Trademark Change - Artwork update | Genx12 Design Transfer |  | Completed | 100 | Ankar Das | false | Medium | 2025-01-21 10:30 | 2025-01-22 |  | 2025-01-24 | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
247 | 2. Trademark Change - Decksheet | Genx12 Design Transfer |  | Completed | 100 | Ankar Das | false | Medium | 2025-01-21 10:30 | 2025-01-22 |  | 2025-01-24 | Ankar Das |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
248 | GenX12 User manual update | Genx12 Design Transfer |  | Completed | 100 |  | false | Medium | 2025-01-21 10:30 | 2025-01-20 |  | 2025-01-20 | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
249 | Genx12 Add COPY functionality in SRS | Genx12 sustenance |  | Completed | 100 | Sindhu Kashyap | false | Medium | 2025-01-21 10:30 | 2025-01-21 |  |  | Sindhu Kashyap | 1/1 | Collect the latest SRS. |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
250 | Working on QARA feedback on DIR and brochure changes | No Bucket |  | Completed | 100 | Sindhu Kashyap | false | Medium | 2025-01-20 10:30 | 2025-01-20 |  |  | Sindhu Kashyap |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
251 | 1. • Our heart rate works on average, so it takes some time to reflect. • But we can discuss and decide on the solution. | GenX12i Series observations from Camp |  | Completed | 100 |  | false | Medium | 2025-01-19 10:30 | 2024-09-07 01:00 |  |  | Anubhav Chaudhary | 0/1 | Internal Discussion Required | EXPLAINED:::: - Our heart rate works on average, so it takes some time to reflect. - But we can discuss and decide on the solution. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
252 | 7. If the patient ID and age are left blank, the report prints them as blank. GENERAL OBSERVATIONS: Negative | GenX12i Series observations from Camp |  | Completed | 100 |  | false | Medium | 2025-01-19 10:30 | 2024-09-07 01:00 |  |  | Anubhav Chaudhary | 0/1 | Solution::Space character is cause for blank space.(Same in 9108 too) |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
253 | 4. For Blood pressure (BP), if a two-digit number are entered without (Eg: 90/120), the default value becomes 250/120. GENERAL OBSERVATIONS: Negative | GenX12i Series observations from Camp |  | Completed | 100 | Mahesh DE | false | Medium | 2025-01-19 10:30 | 2024-09-07 01:00 |  |  | Anubhav Chaudhary | 0/1 | Internal Discussion Required | Three numbers can be entered in both systolic and diastolic fields so if we enter 90 in machine, then it treats it as 900 and replaces it with max limit of 250. So, to enter 90 we shall enter 090. We can have detailed discussion on this and can conclude. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
254 | Genx3 Print format scanned images updates in user Manual | Genx3 Sustenance |  | Completed | 100 | Sindhu Kashyap | false | Medium | 2025-01-15 10:30 | 2025-01-16 |  | 2025-01-17 | Sindhu Kashyap |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
255 | Sanjana : Setting ALL Ubuntu system ready for development | Feature Addition |  | Completed | 100 |  | false | Medium | 2025-01-15 10:30 | 2024-08-01 01:00 |  |  | Anubhav Chaudhary | 1/6 | Install Cutecom on ubuntu on virtual box; Make Laptop Dual Boot;Install Cutecom And Eclipse;Install ubuntu on virtual box;Install Virtual Box;Install Eclipse on Virtual box System | Activity Put on HOLD as lot of dependency is there and no longer needed. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
256 | PDF Report comparison between old and new Dot Net Board S/W in Genx3 and then work on Genx12 | Bug Fix |  | Completed | 100 | Mahesh DE | false | Medium | 2025-01-15 10:30 | 2024-11-04 |  |  | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
257 | Simulaton:In some reports, the patient's name was missing. | Bug Fix |  | Completed | 100 | Former member;chandan kumar | false | Medium | 2025-01-15 10:30 | 2024-07-23 01:00 |  |  | Anubhav Chaudhary | 0/6 | Executing;Initiating;Simulation Needed(Rare Issue);Planning;Closing;Testing |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
258 | PART1 BUGS 0.81 BTR-470-VER-01-005_0.81v_GENX12i+Latest Report attached | Bug Fix |  | Completed | 100 | Anubhav Chaudhary | false | Medium | 2025-01-15 10:30 | 2024-09-02 01:00 |  |  | Anubhav Chaudhary | 13/19 | 10 LCD_SWEEP_SPEED_ADJUSTMENT_2024;9 Implemented like this (Medians will have additional QRS for Higher HR);18. BUG_LL_MODE_FREQ_MISMATCH_IN_ROW2 30/08/2024;19. FIXED PGSV_PREVIEW_DATA_MISMATCH(04/09/2024);4 FIXED: Trigger_MEM_FULL_MODE_DOES_NOT_CHANGE;2 Major;13 Action Needed (Minor);16 Simulation Needed (Minor);3 Fixed BUG_PRINT_FMT_RHY_DISP_IN_SNAP_SHOT_VIEW;17 Simulation Needed (Major Rare);14 Action Needed BLOCK_AUTO_ALL_LEAD_GAIN_IN_freeze ;8 Simulation Needed (Major);7 Simulation Needed (Major) AUTO_SHUTDOWN_IN_TRIGGER_MODE_BUG ;1 FIXED;20. 03_09_2024 LL_MODE_10_5_FIX;6 Action needed(PC) (Will be Done in next );15 Simulated Needed (Major) BUG_LL_MODE_FREQ_MISMATCH_IN_ROW2 (Tested on 13/09/2024;12 Simulation Needed (Major) BUG_LL_MODE_FREQ_MISMATCH_IN_ROW2 (Tested on 13/09/2024);5 Action needed(Ping) (Will be Done in next ) | **As on 6 OCT 2024 7PM** **Total 8 left2 WE WILL NOT DO.6 left** **As on 13 Sep 2024 7PM** **Total 11 leftMinor 2Major 8 , 1 criticalIn Progress 2(22,25)** **As on 13 Sep 2024 4PM** Total 13 left Minor 2 Major 10 , 1 critical In Progress 2(22,25) So after fixing in progress one Total 11 left(10 major 1 minor). As on 4 Sep 2024 • 7 fixed bug. • 10 simulated bugs. • 6 bugs which could not be simulated. • 3 Random occurrences. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
259 | PART2 BTR-470-VER-01-005_0.81v_GENX12i+ | Bug Fix |  | Completed | 100 | Anubhav Chaudhary | false | Medium | 2025-01-15 10:30 | 2024-09-04 01:00 |  |  | Anubhav Chaudhary | 6/6 | 24 ////MOTOR_MOVES_AFTER_PAPER_OUT_BUG (12-09-2024);25 Simulation (In Progress) FREEZE_SAVE_AND_PRINT_TERMINATE_WITH_PAPER_SENSOR_ON_HOME_SCREEN_MODE_CH;23 FIXED BL Restore added;21 Simulation needed MAJOR (Name changed to 65);26 FIXED in 0.82(This got fixed with other Bugs);22 Wrong Pat ID Search(Minor) (In progress) |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
260 | Simulation: In one case, the printed heart rate (HR) differed from the HR displayed on the screen. | Bug Fix |  | Completed | 100 | Former member;chandan kumar | false | Medium | 2025-01-15 10:30 | 2024-07-23 01:00 |  |  | Anubhav Chaudhary | 0/1 | Simulation Needed(Rare Issue) | rare case and both HR are computed on different scenario |  | Item | sites/ECGTeam761/Lists/Genx sustenance
261 | BTR-470-VER-01-005- | Bug Fix |  | Completed | 100 | Anubhav Chaudhary | false | Medium | 2025-01-15 10:30 | 2024-08-12 01:00 | 2024-08-12 01:00 | 2024-09-06 01:00 | Anubhav Chaudhary | 0/4 | If we operate machine in Page Save or Auto mode with “Trigger Sample” -;then status on screen does not shows current mode even if we press mode key.;Data not transferring through USB in HL7 format in case “QRS NOT DETECTED”;Bug: If we operate machine in Page Save or Auto mode with “Trigger Sample” and Memory full setup is | **BTR-470-VER-01-005-This bug report has some Issues - working on these issues ** **** **3.Machine stuck in Loop of transferring the same data again and in PAGESAVE TRIGGER SAMPLE mode if records stored are 399 and “Auto Export to USB” isenabled** **** **4.Two gain marks are coming on second page when interpretation is set asdetailed.** **** **5.ECG Graph on screen is inaccurate as per Heartrate and as per thermalprint. E.g. for 30BPM HR 10 big boxes should be covered, but on screen it iscovering 11 boxes, similarly for 60BPM it should cover 5 boxes but it iscovering 5 and a half boxes** **** **6.Unit shuts down by showing “AUTO SHUTDOWN INITIATED” when machine is in“TRIGGER SAMPLE” mode and “AUTO EXPORT TO USB” operation going on** **** **** **** **¦ <ol><li><span><b><span lang="EN-US">4.&nbsp; Bug:</span></b><span lang="EN-US"> </span><span style="font-size:10pt;" lang="EN-US">If we operate machine in Page Save or Auto mode with “Trigger Sample” and Memory full setup is notify, then status on screen does not shows current mode even if we press mode key.&nbsp; &nbsp; &nbsp; &nbsp; &nbsp; &nbsp; &nbsp; &nbsp; &nbsp; &nbsp; &nbsp; &nbsp; &nbsp; &nbsp; &nbsp; &nbsp; &nbsp; &nbsp; &nbsp;Fixed: Trigger_MEM_FULL_MODE_DOES_NOT_CHANGE</span></span></li></ol> ¦¦ --- ¦** |  | Item | sites/ECGTeam761/Lists/Genx sustenance
262 | Testing and verification Genx12(0.81) | To do |  | Completed | 100 | Former member | false | Medium | 2025-01-15 10:30 | 2024-02-02 |  |  | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
263 | RAM test for using it to its full capability and Heap overlap. | Bug Fix |  | Completed | 100 | Someshwaran S | false | Medium | 2025-01-15 10:30 | 2024-11-05 |  |  | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
264 | Genx12 Print format scanned images updates in user Manual --- By Sindhu | Genx12 Design Transfer |  | Completed | 100 | Sindhu Kashyap | false | Important | 2025-01-15 10:30 | 2025-01-11 | 2025-01-13 | 2025-01-15 | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
265 | Genx3- record Transfer Through FTP | Genx3 Sustenance |  | Completed | 100 | Mahesh DE | false | Medium | 2025-01-15 10:30 | 2025-01-06 | 2025-01-06 | 2025-01-24 | Mahesh DE | 1/2 | mark all record - not working;Genx3- FTP - Transfer PDF ,HL7,and BPL format | Record transfer - to FTP server at version 3.04 Long lead |  | Item | sites/ECGTeam761/Lists/Genx sustenance
266 | PRR and Software anomalies | Genx12 15 Jan Release |  | Completed | 100 | Someshwaran S | false | Medium | 2025-01-12 10:30 | 2025-01-13 |  | 2025-01-13 | Someshwaran S |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
267 | Impedance values noted for ITC & Tech PCB's (VV15) | Vivid Vue |  | Completed | 100 | samiksha shetty | false | Medium | 2025-01-12 10:30 | 2025-01-13 |  |  | samiksha shetty |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
268 | Visual inspection of the components-mounted/non-mounted on the PCB's (VV15) | Vivid Vue |  | Completed | 100 | samiksha shetty | false | Medium | 2025-01-12 10:30 | 2025-01-13 |  |  | samiksha shetty |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
269 | Rework for 10 PCB's (Keypad, ITC, etc) for assembling 10 units (VV15) | Vivid Vue |  | Completed | 100 | samiksha shetty | false | Medium | 2025-01-12 10:30 | 2025-01-13 |  |  | samiksha shetty |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
270 | Rework for 25 (LCD)Display cables (VV15) | Vivid Vue |  | Completed | 100 | samiksha shetty | false | Medium | 2025-01-12 10:30 | 2025-01-13 |  |  | samiksha shetty |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
271 | Noting the NIBP values at different angles of the NIBP module (VV15) | Vivid Vue |  | Completed | 100 | samiksha shetty | false | Medium | 2025-01-12 10:30 | 2025-01-13 |  |  | samiksha shetty |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
272 | ECN Handle for Genx3 | To do |  | Completed | 100 |  | false | Medium | 2025-01-12 10:30 | 2024-02-02 |  |  | Anubhav Chaudhary |  |  | Done By ECG TEAM |  | Item | sites/ECGTeam761/Lists/Genx sustenance
273 | Verification and Validation Report for Genx12 machine for Volvo Camp with all data Backup. | To do |  | Completed | 100 | Former member | false | Medium | 2025-01-12 10:30 | 2024-07-16 01:00 |  |  | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
274 | Reviewer for GenX12 mentioned document-Beryl | To do |  | Completed | 100 | beryl paulson | false | Medium | 2025-01-12 10:30 | 2024-07-29 01:00 | 2024-07-04 01:00 | 2024-09-05 01:00 | beryl paulson | 13/13 | 6.Product Brochure-28-8-2024;1.User manual -2-9-2024;3.Report on Clinical evaluation- 31-8-2024 ;2.TechnicalSpecifications-30-7-2024 -;9.Product_Life Declaration -3-9-2024 - ;13.Service manual -5-8-2024;11.Manufacturing and Test License -1-9-204;8.Environmental and Reliability test -1-9-2024;4.Risk management file- 29-8-2024;10.Component Approval Note -1-9-2024;5.Design transfer checklist -2-9-2024;7.Bio Compatibility Test Report -1-9-2024;12.Product training 5-8-2024 | Dependencies:- Mechanical 1.User manual - completed (Changed according to new spare part list) 2.TechnicalSpecifications- 30-7-2024 - Completed 3.Report on Clinical evaluation- 31-8-2024 Need to review 4.Risk management file- 29-8-2024 in progress 5.Design transfer checklist -2-9-2024 in progress 6.Product Brochure-28-8-2024 Completed 7.Bio Compatibility Test Report -1-9-2024 Completed 8.Environmental and Reliability test -1-9-2024 completed 9.Product_Life Declaration -3-9-2024 - Completed 10.Component Approval Note -1-9-2024 Completed 11.Manufacturing and Test License -1-9-204 Completed 12.Service manual -27-8-2024 completed |  | Item | sites/ECGTeam761/Lists/Genx sustenance
275 | Training Minutes to Discuss with Inder Sir | To do |  | Completed | 100 | Anubhav Chaudhary;Someshwaran S | false | Medium | 2025-01-12 10:30 | 2024-09-11 01:00 |  | 2024-09-12 01:00 | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
276 | (Sanjana)Interpretation and Measurement Matrix (***) assessment and | To do |  | Completed | 100 |  | false | Medium | 2025-01-12 10:30 | 2024-11-04 |  |  | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
277 | (Sanjana)Border line ECG assessment and testing. | To do |  | Completed | 100 |  | false | Medium | 2025-01-12 10:30 | 2024-11-04 |  |  | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
278 | Long Lead ECR ECN Documents | Genx3 Sustenance |  | Completed | 100 | Someshwaran S | false | Medium | 2025-01-10 10:30 | 2025-01-11 |  | 2025-01-10 | Someshwaran S |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
279 | Concurrence With Sankar Regarding | Genx12 15 Jan Release |  | Completed | 100 |  | false | Medium | 2025-01-09 10:30 | 2025-01-09 |  |  | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
280 | ECG record Gain testing | Genx12 15 Jan Release |  | Completed | 100 | Someshwaran S | false | Medium | 2025-01-09 10:30 | 2025-01-06 |  | 2025-01-10 | Anubhav Chaudhary | 8/9 | Home screen gain should not have any impact on ECG rec gain;check the trace in snap shot and scroll view on LCD ,has to be same as ECG rec gain;If any update happens ,need to test the same ;Check the Median Gain on second Page;If ECG rec gain is off,then the saved /set gain value will be appeared on above cases;check ECG Rec pendrive transfer,FTP,External usb printer ecg record gain and trace.;Check the gain value and gain mark symbol in snap ,scroll view record print same as ECG rec.gain;Check the trace in prints(ECG record snap shot,scroll view prints) it has to same as ECG rec.gain;In ECG Records print, Check the Gain Mark and gain value ,it has to same as ECG record gain |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
281 | ECG record Gain Implementation | Genx12 15 Jan Release |  | Completed | 100 | Anubhav Chaudhary | false | Medium | 2025-01-09 10:30 | 2025-01-06 |  | 2025-01-10 | Anubhav Chaudhary | 16/17 | Median Gain Implementation on second Page;NIC Board transfer gain Impact(SPI).;In ECG records snap shot and scroll view the gain value and trace should be implemented on LCD;In ECG records print on Scroll, snap shot view Trace and gain value same as ECG reocrd gain .;Touch Addition to ECG Record gain UI Menu;UI Addition of option for ECG rec gain and Save to EEPROM;COPY PRINT SHALL HAVE ecg REC GAIN ;From ECG Rec pendrive transfer,FTP,External usb printer ecg record gain and trace.;Print from ECG Record Scroll view screen -Gain Symbol, Gain value and ECG Trace.;Print from ECG Record view screen -Gain Symbol, Gain value and ECG Trace.;View on LCD from ECG Record Scroll view screen -Gain Symbol, Gain value and ECG Trace.;Live print gain shall have no impact on ECG Rec gain.;View on LCD from ECG Record SNAP SHOT screen -Gain Symbol, Gain value and ECG Trace.;For LIVE rec(Auto FTP and USB) pendrive transfer,FTP,External usb printer ecg record gain and trace.;View on LCD from ECG Record view screen -Gain Symbol, Gain value and ECG Trace.;Print from ECG Record SNAP SHOT screen -Gain Symbol, Gain value and ECG Trace.;If any input from chandan ,fix the same . |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
282 | NOENATAL ACCESSORIES Feasibility. | Feature Addition |  | Completed | 100 | samiksha shetty | false | Medium | 2025-01-09 10:30 | 2024-07-29 01:00 | 2024-08-23 01:00 |  | samiksha shetty | 2/7 | Source the sample.;Do atleast 10 live tests.;Discuss with Inder Sir and SCM team for neonatal electrode. ;test the samples and compare with predicate device.;Test on simulator;Send the electrodes for validation.;Get the validation report from hospital or diagnostic center.(product team will help on this) | Dependencies:- 1. Sourcing the electrodes. 2. Sales team for the validation report. 3. Validation timeline/duration. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
283 | Update the Export Print Format to 3X4,3x4+1R,3X4+3R,6x2,6x2+1R.Currently we provide only 12X1. | To do |  | Completed | 100 | Mahesh DE | false | Medium | 2025-01-05 10:30 | 2024-07-29 01:00 |  |  | Mahesh DE | 2/6 | Test 6x2 Format if its working correctly;Implement 6x2;Implement 3x4 + 3R;Implement 6x2+1 R;Implement 3x4;Understand ECG plot in Dot Net Code. | Update the Export Print Format to 3X4,3x4+1R,3X4+3R,6x2,6x2+1R.Currently we provide only 12X1. for Genx12 |  | Item | sites/ECGTeam761/Lists/Genx sustenance
284 | Pacemaker mark is deformed in page save print Occurrence : 2/20 | Genx12 15 Jan Release |  | Completed | 100 | Mahesh DE | false | Medium | 2025-01-05 10:30 | 2025-01-06 | 2025-01-06 | 2025-01-06 | Mahesh DE | 3/3 | change the width of Pacer Pulse- "u16PacerPulsePrintCounter>10" made as 24;connect and check the Pacemaker ;" application.c" change the space as u16PacerPulsePrintCounter>30: | connect and check the Pacemaker - few records showing as Spike on edge of Pacer Pulse. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
285 | GENX3 LL Mode ECRDocuments | Genx3 Sustenance |  | Completed | 100 | Someshwaran S | false | Urgent | 2025-01-05 10:30 | 2025-01-06 | 2024-12-30 | 2025-01-06 | Someshwaran S | 5/5 | Create FMI,DVP,DVR Feedback form;Check the impact on documents;Fill the Gaps Observed in genx3 /genx3i;Update as per Chandrasekhar earlier comments;create Latest DRR |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
286 | 9. Patient Name : 65 by default when we press START/STOP button after 1-2 secs ECG waveform printing | Genx12 15 Jan Release |  | Completed | 100 | Anubhav Chaudhary | false | Medium | 2025-01-05 10:30 | 2025-01-06 |  | 2024-12-30 | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
287 | 16 .Patient record with V2 lead off is getting saved into memory and if auto USB transfer is enabled then machine/keypad becomes unresponsive after transferring record. | Genx12 15 Jan Release |  | Completed | 100 | Anubhav Chaudhary | false | Medium | 2025-01-05 10:30 | 2025-01-06 |  | 2025-01-02 | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
288 | Settling Time in ECG Board was Observed. It was fixed by reducing delay in ECG Board. | Genx12 15 Jan Release |  | Completed | 100 | Anubhav Chaudhary | false | Medium | 2025-01-05 10:30 | 2025-01-06 |  | 2024-12-20 | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
289 | 3. FTP_TRANSFER_FAIL_FIX | Genx12 15 Jan Release |  | Completed | 100 | Anubhav Chaudhary | false | Medium | 2025-01-05 10:30 | 2025-01-06 |  | 2024-12-30 | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
290 | external Printer -for Rhythm Lead | Genx12 15 Jan Release |  | Completed | 100 | Mahesh DE | false | Medium | 2025-01-05 10:30 | 2025-01-06 | 2024-12-17 | 2024-12-17 | Mahesh DE | 3/3 | made few changes in the "App configuration.c" file related to Print format;check Rhythm lead- and compare the PDF format ;changes in -u16FormatSelectedVal Variable | As per Rhythm lead- the PDF format is printing |  | Item | sites/ECGTeam761/Lists/Genx sustenance
291 | Genx3- Record Transfer | Genx3 Sustenance |  | Completed | 100 | Mahesh DE | false | Medium | 2025-01-05 10:30 | 2025-01-06 | 2025-01-10 |  | Mahesh DE | 3/3 | Record transfer to USB- HL7 format;Record Trnsfer to USB - BPL format;Record Transfers to USB - PDF format | the storage of Genx3 is 500 record so all record should transfer to USB |  | Item | sites/ECGTeam761/Lists/Genx sustenance
292 | 23 NIBP | Genx12 15 Jan Release |  | Completed | 100 | Mahesh DE | false | Medium | 2025-01-05 10:30 | 2025-01-06 | 2024-12-25 | 2024-12-30 | Mahesh DE | 5/5 | if it is Limit - store the value in Patient info - and Update the Variable;And after 2 sec Disappear ;Check the AGE limit ,Height and Weight Limit;check the existing Limit- NIBP ;Create CAUTION popup massage- if not in limit -shows the CAUTION Message and apply the Limit. | NIBP- the few Feature option is added in the Patient info Option at main menu screen if Systolic and Dystalic not in Range - it shows the POP UP MESSGE as Its is not in Range , |  | Item | sites/ECGTeam761/Lists/Genx sustenance
293 | Alternate components for GenX12 | Obselence |  | Completed | 100 | nikhil siby | false | Medium | 2024-12-01 10:30 | 2024-07-24 01:00 | 2024-08-05 01:00 | 2024-09-11 01:00 | nikhil siby | 8/8 | BAV99LT1G;1N4148WS-E3-08;IRF9388TRPBF;MT41K256M16TW-107:P;PESD5V0X1BL,315;MSASU168SB7224KTNB25;SN74LVC1G08IDCKRQ1;Report | SN74CBTLV3257PWR, SN74LV4052APWR can be considered in the next update since some layout changes are requires. feasibility study also need to do. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
294 | PC Communications stops in between for all Genx Devices | Bug Fix |  | Completed | 100 | samiksha shetty | false | Medium | 2024-11-27 10:30 | 2024-08-19 01:00 |  |  | samiksha shetty | 0/2 | Setup RT View software in desktop(Take Ajay's help);Start communication with PC and Observe |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
295 | GRI Summary statement Addition Testing | To do |  | Completed | 100 | samiksha shetty | false | Medium | 2024-11-10 10:30 | 2024-11-04 |  |  | samiksha shetty |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
296 | Set Genx3 ECN Design Review Meeting(Monday) | To do |  | Completed | 100 | Anubhav Chaudhary | false | Important | 2024-11-05 10:30 | 2024-08-22 01:00 |  | 2024-08-30 01:00 | Anubhav Chaudhary | 1/3 | Review will be done;DRR is created by Mahesh;Meeting set and Done |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
297 | (Sanjana)GRI Averaging Correction code update to polynomial | To do |  | Completed | 100 |  | false | Medium | 2024-11-05 10:30 | 2024-11-04 |  |  | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
298 | MARK_ALL_WITH_LL_REC_STATUS_UPDATE_BUG | Bug Fix |  | Completed | 100 | Anubhav Chaudhary | false | Medium | 2024-11-05 10:30 | 2024-11-04 |  | 2024-11-06 | Anubhav Chaudhary | 2/2 | Implementation Done;Developer testing Done |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
299 | 11 Simulation Needed (Major) PACS | Bug Fix |  | Completed | 100 | Anubhav Chaudhary | false | Important | 2024-11-03 10:30 | 2024-09-23 01:00 |  |  | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
300 | ////MOTOR_MOVES_AFTER_PAPER_OUT_BUG | Bug Fix |  | Completed | 100 |  | false | Medium | 2024-11-03 10:30 | 2024-09-12 01:00 |  |  | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
301 | Sanjana Unit test List | Bug Fix |  | Completed | 100 |  | false | Medium | 2024-11-03 10:30 | 2024-09-12 01:00 |  |  | Anubhav Chaudhary | 0/1 | 2nd COPY press, paper sensor ON. Int Detailed, |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
302 | 4.2 Copy not working as per recommendation Genx12 – Gujarat (Ahmedabad) | GenX12i Series observations from Camp |  | Completed | 100 | Anubhav Chaudhary | false | Medium | 2024-11-03 10:30 | 2024-09-07 01:00 |  |  | Anubhav Chaudhary | 2/3 | Need Simulation ;Same as 2,3;Need Discussion |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
303 | 4.5 Printer Ink quality is poor Genx12 – Gujarat (Ahmedabad) | GenX12i Series observations from Camp |  | Completed | 100 | Someshwaran S;nikhil siby | false | Medium | 2024-11-03 10:30 | 2024-09-07 01:00 |  |  | Anubhav Chaudhary | 1/2 | Nikhil, Somesh will check this unit for printing ;Waiting For Machine | - Please ensure that the print head is clean before printing. - It is recommended to use normal print mode which results in longer printer life. After head clean, it was observed and showcased to Shankar and team that the print quality improved. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
304 | Suggestion: HR refresh on Lead Off | GenX12i Series observations from Camp |  | Completed | 100 | Anubhav Chaudhary | false | Medium | 2024-11-03 10:30 | 2024-09-16 01:00 | 2024-09-19 01:00 | 2024-09-20 01:00 | Anubhav Chaudhary | 1/2 | Testing required ;Implemented |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
305 | Samiksha: Story Board | Feature Addition |  | Completed | 100 | samiksha shetty;Ankar Das;Nitin Chavan | false | Medium | 2024-10-23 11:30 | 2024-08-01 01:00 | 2024-08-12 01:00 | 2024-10-24 01:00 | Nitin Chavan | 4/10 | Email to shankar - 23/08/2024;Thorough understanding on usability. ;GUI review with cross functional teams.;Competitor Analysis (ECG models);Come with rough workflows for new machine.;Come with rough GUI for new machine.;Check all ECG machine for GUI.;Study of user manual.;Working with Graphic Designer on various GUI approach based on workflows.;Preparing Story board. | Samiksha is the task owner. Dependencies:- 1. STD 62366 understanding. 2. Procurement of competitor machines. 3. Understanding on the usability of the devices. 4. Availability of the graphic designer. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
306 | 2.3 The copy function is not working as intended GenX12i+ – YKJA4G1002 | GenX12i Series observations from Camp |  | Completed | 100 | Anubhav Chaudhary;chandan kumar | false | Medium | 2024-10-02 11:30 | 2024-09-07 01:00 |  |  | Anubhav Chaudhary | 2/3 | Discussion needed with product Team;(It worked fine)Tried simulation with product team ;Simulation sequence required | - As per our understanding copy function will just print the last saved record after machine boot and if nothing is saved after machine boot then it will not work. Please provide us details of how it is functioning. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
307 | Provide single page PDF and USB Print Report with interpretation. | Feature Addition |  | Completed | 100 | Anubhav Chaudhary;samiksha shetty;Mahesh DE | false | Medium | 2024-10-02 11:30 | 2024-07-29 01:00 | 2024-09-23 01:00 | 2024-09-27 01:00 | samiksha shetty |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
308 | Suggestion: Make Default Gain as 10/5. | GenX12i Series observations from Camp |  | Completed | 100 | Anubhav Chaudhary | false | Medium | 2024-09-18 11:30 | 2024-09-17 01:00 | 2024-09-19 01:00 | 2024-09-20 01:00 | Anubhav Chaudhary | 2/2 | Internal testing Done ;Implemented |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
309 | One FG Unit it was observed that in higher HR (>180) GRI was not computed YKJA4G1002 | GenX12i Series observations from Camp |  | Completed | 100 | Anubhav Chaudhary | false | Medium | 2024-09-16 11:30 | 2024-09-17 01:00 |  | 2024-09-17 01:00 | Anubhav Chaudhary | 0/1 | PAPER_SENSOR_ON_GRI_NOT_COMPUTED_IN_SOME_PRT_FMTS | PAPER_SENSOR_ON_GRI_NOT_COMPUTED_IN_SOME_PRT_FMTS |  | Item | sites/ECGTeam761/Lists/Genx sustenance
310 | COPY Feature to add in user Manual. add | Genx12 Design Transfer |  | Completed | 100 | samiksha shetty | false | Medium | 2024-09-15 11:30 | 2024-09-11 01:00 |  |  | samiksha shetty |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
311 | Genx3 ECR related Docs, Software Release notes ,Software Modification Document ,Software Problem Resolution Report, Internal release notes | To do |  | Completed | 100 | Mahesh DE | false | Medium | 2024-09-12 11:30 | 2024-07-09 01:00 | 2024-07-11 01:00 | 2024-08-22 01:00 | Mahesh DE | 7/7 | Software Release notes;Update RMR;Genx3 ECR related Docs;Update RMP;Internal release notes ;Software Problem Resolution Report;Software Modification Document | RMR and RMP not applicable Software Modification Documents -with version 379-SM-01 with rev No-1.7 updated for Genx3 1. Software problem resolution Report- "379_PRR_01 REV 1-7" Updated |  | Item | sites/ECGTeam761/Lists/Genx sustenance
312 | Freeze F2 option-> doesn't not display abort msg INVALID ID | Bug Fix |  | Completed | 100 |  | false | Medium | 2024-09-11 11:30 | 2024-09-12 01:00 | 2024-09-12 01:00 | 2024-09-12 01:00 | Anubhav Chaudhary | 1/1 | //ADDED WITHOUT_NAME_ID_FREEZE_F1_DOES_NOT_DISPLAY_MSG_STR_PLEASE_ENTER_VALID_NAME_OR_ID (12/09/2024 | //ADDED WITHOUT_NAME_ID_FREEZE_F1_DOES_NOT_DISPLAY_MSG_STR_PLEASE_ENTER_VALID_NAME_OR_ID (12/09/2024) |  | Item | sites/ECGTeam761/Lists/Genx sustenance
313 | Create Questionnaire Google Form for Service Training | To do |  | Completed | 100 | samiksha shetty | false | Medium | 2024-09-10 11:30 | 2024-09-11 01:00 |  | 2024-09-11 01:00 | samiksha shetty |  |  | Take input from old question. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
314 | Unit Monkey testing to start and Discuss 0.80 issues with product team | No Bucket |  | Completed | 100 | Anubhav Chaudhary | false | Medium | 2024-09-10 11:30 | 2024-08-29 01:00 |  | 2024-09-09 01:00 | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
315 | 3. The saved record has a gain of 10/5, but when exported to PDF and viewed, it shows a gain of 5 mm/mV. GENERAL OBSERVATIONS: Negative | GenX12i Series observations from Camp |  | Completed | 100 | Anubhav Chaudhary | false | Medium | 2024-09-10 11:30 | 2024-09-07 01:00 |  | 2024-09-11 01:00 | Anubhav Chaudhary | 2/3 | 2.5 Gain from volvo F/B will be resolved with this;External Discussion required;Internal Discussion Required | As of now PDF report does not support 10/5 gain. So as immediate solution we can make it as 10 instead of 5. /////////FOR_10_5_MAKE_DEFAULT_AS_10 //11/09/2024 The saved record has a gain of 10/5, but when exported to PDF and viewed, it shows a gain of 5 mm/mV. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
316 | 5. Limb leads signals have more interference compared to chest leads. GENERAL OBSERVATIONS: Negative | GenX12i Series observations from Camp |  | Completed | 100 |  | false | Medium | 2024-09-09 11:30 | 2024-09-07 01:00 |  |  | Anubhav Chaudhary | 1/1 | External Discussion | - If electrodes are put on the joints, it adds tremor to the ECG leads - Need more clarity on the type of noise. - If images or printouts are provided, it will be helpful. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
317 | 1.1 The display waveform/trace was not available in specific formats, such as 6X1 and 6X2, although these formats were visible in printouts. GenX12i –- XQJA4G1002 | GenX12i Series observations from Camp |  | Completed | 100 | Anubhav Chaudhary | false | Medium | 2024-09-09 11:30 | 2024-09-07 01:00 |  | 2024-09-11 01:00 | Anubhav Chaudhary | 0/5 | Implement Fix;Integration Test;Observation and Analysis ;Plan for Fix;Unit Test |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
318 | 4.4 Footer print is mis-aligned though paper is proper. Genx12 – Gujarat (Ahmedabad) | GenX12i Series observations from Camp |  | Completed | 100 |  | false | Medium | 2024-09-09 11:30 | 2024-09-07 01:00 |  |  | Anubhav Chaudhary | 0/1 | Waiting for machine. | Correct spacer position should be ensured before printing. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
319 | 3.1 The display waveform/trace is not available in specific formats like 6X1 and 6X2, although these formats are visible in printouts. GenX12i – XQJA4G1006 | GenX12i Series observations from Camp |  | Completed | 100 | Anubhav Chaudhary;chandan kumar | false | Medium | 2024-09-09 11:30 | 2024-09-07 01:00 |  |  | Anubhav Chaudhary |  |  | - It seems as the same issue as mentioned in 1.1. - We have received this unit and working on it. Correct us if our observation is wrong. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
320 | 3.2 The machine freezes in Auto mode, resulting in a blank screen with no waveforms displayed. GenX12i – XQJA4G1006 | GenX12i Series observations from Camp |  | Completed | 100 | Anubhav Chaudhary;chandan kumar | false | Medium | 2024-09-09 11:30 | 2024-09-07 01:00 |  |  | Anubhav Chaudhary | 0/1 | Need Discussion (Check if it is related to 1.1) | - It seems as the same issue as mentioned in 1.3. - We have received this unit and working on it. Correct us if our observation is wrong. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
321 | Reviewer for GenX12 mentioned Document-NIKHIL | To do |  | Completed | 100 | nikhil siby | false | Medium | 2024-09-08 11:30 | 2024-07-29 01:00 | 2024-07-29 01:00 | 2024-09-05 01:00 | nikhil siby | 8/8 | 2.Unit_Alignment_Procedure -5-08-2024;1.Board Alignment_Procedure -5-08-2024;6.New electrical component drawing datasheet 3-9-2024;8.Service manual -5-8-2024;4.Artworks & Labels-07-08-2024;7.Gerber -4-9-2024;5.Schematic updation -27-08-2024;3.Inspection_Procedure-6-08-2024 | 1.Board Alignment_Procedure -5-8-2024 -completed before time(1-8-2024) 2.Unit Alignment_Procedure -5-08-2024 -Completed before time (1-8-2024) 3.Inspection_Procedure-6-08-2024 - Completed before time (5-8-2024) 4.Artworks & Labels- 29-08-2024 Completed 5.Schematic -27-08-2024 Completed 6.New electrical component drawing datasheet -3-9-2024 7.Gerber -4-9-2024 8.Service manual -29-8-24 Needs Update |  | Item | sites/ECGTeam761/Lists/Genx sustenance
322 | 5.1 Print quality low. Without dark trace the print is very much faded. Genx12 – Random checkup | GenX12i Series observations from Camp |  | Completed | 100 |  | false | Medium | 2024-09-06 11:30 | 2024-09-07 01:00 |  |  | Anubhav Chaudhary |  |  | - Recheck the printer assembly for loose fitment. - Please ensure that the print head is clean before printing. - It is recommended to use normal print mode which results in longer printer life. After head clean, it was observed and showcased to Shankar and team that the print quality improved. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
323 | 5.5 Baseline shift to be looked into as the chest leads are shifted down the page. Genx12 – Random checkup | GenX12i Series observations from Camp |  | Completed | 100 |  | false | Medium | 2024-09-06 11:30 | 2024-09-07 01:00 |  |  | Anubhav Chaudhary |  |  | - This issue is fixed in software version 0.82. Note: As per our observation we have fixed it. If simulation steps are provided, we will re-verify. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
324 | 6. Data export to HL7 aborts midway, and once HL7 export fails, subsequent exports to PDF and BIN formats also fail. GENERAL OBSERVATIONS: Negative | GenX12i Series observations from Camp |  | Completed | 100 |  | false | Medium | 2024-09-06 11:30 | 2024-09-07 01:00 |  |  | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
325 | 5.4 Without pressing the print button there is printing initiated with a fully dark coloured trace in L-L mode. Sequence followed: Mode -> L-L mode -> Print -> Print ended -> without pressing PRINT button, print is initiated. Genx12 – Random checkup | GenX12i Series observations from Camp |  | Completed | 100 | Anubhav Chaudhary | false | Medium | 2024-09-06 11:30 | 2024-09-07 01:00 |  |  | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
326 | 5.3 In the end of every L-L mode print there is a coloured trace Genx12 – Random checkup | GenX12i Series observations from Camp |  | Completed | 100 |  | false | Medium | 2024-09-06 11:30 | 2024-09-07 01:00 |  |  | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
327 | 5.2 Straight lines observed in L-L mode while printing. Genx12 – Random checkup | GenX12i Series observations from Camp |  | Completed | 100 | Anubhav Chaudhary | false | Medium | 2024-09-06 11:30 | 2024-09-07 01:00 |  |  | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
328 | Reviewer for GenX12 mentioned Document -Somesh | To do |  | Completed | 100 | Someshwaran S | false | Medium | 2024-09-06 11:30 | 2024-07-29 01:00 | 2024-07-29 01:00 | 2024-08-30 01:00 | Someshwaran S | 7/7 | 2.Design_Parts_list 3-09-2024;3.Assembly Drawings ( somesh ,jeswin) 29-8-2024;1.wiring Diagram-29-7-2024 -DONE;6.Assembly Tree Diagram-31-7-2024 -DONE;7.Cable Drawings -1-08-2024;5.Critical Component List -30-7-2024 -DONE;4.Mechanical Drawings (Somesh ,jeswin) 29-8-2024 | Reviewer for these document-Somesh 1.wiring Diagram-31-7-2024 -DONE -Completed before time(30-7-2024) 2.Design_Parts_list 3.Assembly Drawings ( somesh ,jeswin) - mechanical needs to give output 4.Mechanical Drawings (Somesh ,jeswin) -mechanical needs to give output 5.Critical Component List -30-7-2024 -DONE 6.Assembly Tree Diagram-31-7-2024 -DONE 7.Cable Drawings -1-08-2024 - completd before time (31-7-2024) |  | Item | sites/ECGTeam761/Lists/Genx sustenance
329 | 4.1 Baseline restore in auto mode (provided in LL mode and manual but no in Auto) Genx12 – Gujarat (Ahmedabad) | GenX12i Series observations from Camp |  | Completed | 100 |  | false | Medium | 2024-09-06 11:30 | 2024-09-07 01:00 |  |  | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
330 | 1.2 Data export through HL7 frequently aborted midway. GenX12i –- XQJA4G1002 | GenX12i Series observations from Camp |  | Completed | 100 |  | false | Medium | 2024-09-06 11:30 | 2024-09-07 01:00 |  |  | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
331 | Genx12 9 inch LCD test report and verification | Obselence |  | Completed | 100 | Someshwaran S | false | Medium | 2024-09-05 11:30 | 2024-08-27 01:00 |  | 2024-09-05 01:00 | Someshwaran S | 3/3 | Comparison and test report.;Comparing With current LCD.;Verify the performance. |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
332 | Genx3 Long Lead- Design transfer | To do |  | Completed | 100 | samiksha shetty | false | Medium | 2024-09-04 11:30 | 2024-08-28 01:00 |  |  | samiksha shetty | 4/4 | Genx3 User Manual approval Note;Genx3 User Manual ;Genx3i User Manual;Genx3i User Manual Approval Note | Take the Previous Version from Server- User Manual and User Manual approval Note Update and send to Beryl Mail - CC should be Anubhav ,Nitin and Mahesh |  | Item | sites/ECGTeam761/Lists/Genx sustenance
333 | Z fold comes out printer bracket gap | Bug Fix |  | Completed | 100 | Someshwaran S | false | Important | 2024-09-04 11:30 | 2024-07-29 01:00 | 2024-08-05 01:00 |  | Someshwaran S | 6/6 | By august 1st week expecting samples for testing. ;Discussed the same with Girish and Mechanical team, they gave one paper stopper as a solution. Drawi;Test the samples on 5 machines ;Need to send to scm to get quote for bulk production;Deploy;Capture the drawing in DOF | Dependencies: Mechanical Team. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
334 | Assemble genx12+ plastics and discuss with Product Team. | Feature Addition |  | Completed | 100 | nikhil siby | false | Medium | 2024-09-02 11:30 | 2024-08-11 01:00 |  | 2024-09-03 01:00 | nikhil siby | 1/1 | Check with Inder Sir for Shaju F/B | Dependencies:- Inder sir Shown to Kanan and Shankar. Escalated to guru Sir. Took Plastics to Factory and shown to Shaju. Shaju will do mold analysis and speak to vendor. Discussed with shaju and the initial feedback will be given within 2-3 days As per the discussion with girish sir existing plastics will be painted. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
335 | ADDED:LL Mode/Manual Mode>Press F4/BL Restore Softkey. | Bug Fix |  | Completed | 100 |  | false | Medium | 2024-09-01 11:30 | 2024-08-30 01:00 | 2024-08-29 01:00 | 2024-08-29 01:00 | Anubhav Chaudhary |  |  | Intentionally it was done. Soft key only resets min and MAX. But BL key will restart ECG Board. We have made it same for soft key too. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
336 | Simulation: The Copy feature is not functioning as expected. | Bug Fix |  | Completed | 100 | Former member;chandan kumar | false | Medium | 2024-09-01 11:30 | 2024-07-23 01:00 |  |  | chandan kumar | 0/1 | Simulation Needed(Rare Issue) | Feature is working as intended. Already shown to Pooja and Deeksha. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
337 | Simulation: In two instances, the Page Save mode saved the previous record instead of the current one. | Bug Fix |  | Completed | 100 | Former member;chandan kumar | false | Medium | 2024-09-01 11:30 | 2024-07-23 01:00 |  |  | chandan kumar | 1/1 | Simulation Needed(Rare Issue) | Simulation steps : Page Save Mode>Records>Open an Older Record in Scroll View>Exit>Page save mode>Press START/STOP for Page Save Preview (Scroll View) |  | Item | sites/ECGTeam761/Lists/Genx sustenance
338 | BUG_LL_MODE_FREQ_MISMATCH_IN_ROW2 30/08/2024 | Bug Fix |  | Completed | 100 | Anubhav Chaudhary | false | Medium | 2024-09-01 11:30 | 2024-08-30 01:00 |  | 2024-08-30 01:00 | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
339 | Menu>settings>page save preview>Snapshot view>Save & Exit>Page save mode>Print Format : 3x4+3Rh, 3x4+1Rh format or 6x2+1Rh >Press START/STOP. In case Snapsho t view comes after “AUTO TRANSFER TO USB” operation then “IInd” lead name is printed in place of | Bug Fix |  | Completed | 100 | Anubhav Chaudhary | false | Medium | 2024-09-01 11:30 | 2024-08-30 01:00 | 2024-08-30 01:00 | 2024-08-30 01:00 | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
340 | MDD to MDR Conversion | To do |  | Completed | 100 | Anubhav Chaudhary;samiksha shetty;beryl paulson;Nitin Chavan | false | Medium | 2024-08-29 11:30 | 2024-08-23 01:00 | 2024-08-27 01:00 |  | beryl paulson | 0/3 | Preparing the technical file;Cordination with testing lab (Carry out testing);Understanding Doc gap | Dependencies: Gap Analysis as per standards. Understanding the Doc format. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
341 | ECG Graph on screen is inaccurate as per Heartrate and as per thermal print. E.g. for 30BPM HR 10 big boxes should be covered, but on screen it is covering 11 boxes, similarly for 60BPM it should cover 5 boxes but it is covering 5 and a half boxes. | Bug Fix |  | Completed | 100 | Anubhav Chaudhary | false | Medium | 2024-08-27 11:30 | 2024-08-28 01:00 | 2024-08-27 01:00 | 2024-08-28 01:00 | Anubhav Chaudhary | 0/1 | LCD_SWEEP_SPEED_ADJUSTMENT_2024 |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
342 | Genx12 SPL in Bpl format | To do |  | Completed | 100 |  | false | Medium | 2024-08-26 11:30 | 2024-08-27 01:00 | 2024-08-27 01:00 | 2024-08-27 01:00 | Someshwaran S |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
343 | Genx12 Spare Parts List | To do |  | Completed | 100 | Someshwaran S | false | Important | 2024-08-25 11:30 | 2024-08-26 01:00 | 2024-08-19 01:00 | 2024-08-27 01:00 | Someshwaran S | 5/5 | Update SAP code and description in SAP.;Change in diescription w.r.to service pov;Discussion with shakarprasad and assigning new codes;Need to create new Spare parts list;Send a updated SPL to shankar Prasad. |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
344 | review of the Inspection and board alignment procedure | To do |  | Completed | 100 | Someshwaran S;nikhil siby;Nitin Chavan | false | Medium | 2024-08-25 11:30 | 2024-07-13 01:00 | 2024-08-26 01:00 | 2024-08-26 01:00 | nikhil siby | 1/1 | Batt Drain Test(As decided in MTER) |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
345 | Add DEXTOCARDIO CONDITION | To do |  | Completed | 100 | Anubhav Chaudhary | false | Medium | 2024-08-24 11:30 | 2024-07-15 01:00 |  |  | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
346 | Feasibility of adding flap to be assessed. | Feature Addition |  | Completed | 100 | Someshwaran S | false | Urgent | 2024-08-22 11:30 | 2024-07-23 01:00 | 2024-08-23 01:00 | 2024-08-23 01:00 | Anubhav Chaudhary | 1/1 | Discussion with Inder Sir and Girish | Discussed With Girish and Girish gave a paper stopper with adhesive to stop paper flow outside. Testing Completed. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
347 | Whith 399 Rec go in Freeze Screen and press F1 key then Back ground color Change BUG | Bug Fix |  | Completed | 100 | Anubhav Chaudhary | false | Medium | 2024-08-22 11:30 | 2024-08-23 01:00 | 2024-08-21 01:00 | 2024-08-22 01:00 | Anubhav Chaudhary |  |  | //22/08/2024 Chandan BTR(BTR-470-VER-01-005_0.81v_GENX12i.doc) point 3 sub part 5 Memory Setup: Notify>Save 399 records>Auto Mode RT Sample>Press Freeze>Press START/STOP>F1 : SAVE AND PRINT 10 SEC ECG>It shows “No space in memory – Delete some records and try again”>preview screen comes with Black background and ECG wave forms not visible. Machine hangs and needs to restart to function |  | Item | sites/ECGTeam761/Lists/Genx sustenance
348 | Genx12 8 Inch LCD replacement testing. | Obselence |  | Completed | 100 | Someshwaran S | false | Medium | 2024-08-22 11:30 | 2024-08-19 01:00 |  |  | Anubhav Chaudhary | 4/5 | Setting the Right clock for 8 inch LCD;Testing LCD performance for permissible clock frequency(EMC);Burning of 1 nos. 8 inch LCD;Burning of 3 nos. 8 inch LCD;Modifying the code for 8 inch LCD. |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
349 | Add Alphabets (except special Characters) to patient ID field. | Feature Addition |  | Completed | 100 | Anubhav Chaudhary | false | Medium | 2024-08-22 11:30 | 2024-07-23 01:00 |  |  | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
350 | Warranty Label Addition | To do |  | Completed | 100 | beryl paulson | false | Important | 2024-08-22 11:30 | 2024-08-02 01:00 |  |  | Anubhav Chaudhary | 0/2 | Add Label in Genx1, Genx3 and Genx12 user manual;Discuss with Shankar and Inder Sir |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
351 | Add Alphabets(except special Characters) to patient ID field. | Feature Addition |  | Completed | 100 | Anubhav Chaudhary | false | Medium | 2024-08-22 11:30 | 2024-07-29 01:00 |  |  | Anubhav Chaudhary | 3/3 | Asses if Nothing will happens to PDF file creation is we add alphabet to ID ;Export and Upload PDF and HL7 files for testing. ;Implement this in code |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
352 | Reminder: DIACOM and VCG MRT email | Feature Addition |  | Completed | 100 | Anubhav Chaudhary | false | Important | 2024-08-05 11:30 | 2024-08-01 01:00 |  | 2024-08-02 01:00 | Anubhav Chaudhary |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
353 | LL Mode Changes the Print Format as 6ch | Genx12 Sustenance |  | Completed | 100 | Mahesh |  | Medium | 2025-04-16 | 2025-04-14 01:00 | 2025-04-15 01:00 |  |  |  | while changing Mode from one mode to another Mode --the LL mode changes the Print format- - check the related function -add the SetConfiguration |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
354 | User Manual -Genx12 | GENX12 ECR-ECN for SMPS |  | Completed | 100 | sanjana | [8 empty] | Add Neonatal cables in user manual of genx12 as optional accessory, If any optional accessory mentioned, add Neontal cable |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
355 | RT-View Compatibility Test Report (means : S/W work across different OS) | Genx3 Sustenance |  | Not started | 100 | Samiksha |  | Medium | [9 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
356 | Meditix Update---Create RT-View 62304 Report for Genx3 as part of MDD->MDR activity | Genx3 Sustenance |  | Completed | 100 | Sanjana | N.A. | Medium |  | 2025-05-09 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
357 | Meditix Update--- RMF - Software related risks (As part of MDD-> MDR) --- RT view | Genx3 Sustenance |  | Completed | 100 | Sanjana ,Anubhav | N.A. | Medium |  | 2025-05-09 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
358 | Adjust (Precise) Genx12 Gain mark on thermal print | Genx12 sustenance |  | Completed | 100 | Anubhav Chaudhary | N.A. | Medium | 2025-05-19 | 2025-05-12 | 2025-05-11 | 2025-05-16 | Anubhav |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
359 | Check -page Number mis matched at all modes and Freeze modes and Record Data- Print Out format | Genx12 sustenance |  | Dependency (Chk Comment Section) | 100 | Mahesh DE | N.A. | Medium |  | 2025-05-14 | 2025-05-14 |  |  |  | 1.check all Print format |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
360 | Genx3 Upload to GIT HUB grom version 3.01 to 3.04 with all commits. | Genx3 Sustenance |  | Completed | 100 | Anubhav | N.A. | Medium |  | 2025-05-14 | 2025-05-14 | 2025-05-16 | [5 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
361 | AC Line Filter-Noise | Genx12 sustenance |  | Completed | 100 | Mahesh DE | N.A. | Medium | 2025-05-14 | 2025-05-13 | 2025-05-13 | 2025-05-14 | Mahesh DE |  | 1 Check the ADF and MUF - OFF condition 50HZ condition and 60 Hz condition 2. check the previous version of ADF and MUF 3.add new Filter application |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
362 | 1st Lead and 3 lead GUI implement | Genx12 sustenance |  | Completed | 100 | mahesh DE | N.A. | Medium | 2025-05-20 | 2025-05-16 | 2025-05-15 |  | mahesh De |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
363 | (Product team May 2025) Make 6X2+1R as default print format. -- Product team Input (May 2025) | Genx12 sustenance |  | Completed | 100 | Sanjana Sampige | N.A. | Medium |  | 2025-05-23 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
364 | (Product team May 2025) Shrinking of the 3-page report to 2 pages is to be done. (Put on HOLD) | Genx12 sustenance |  | Dependency (Chk Comment Section) | 100 |  | N.A. |  |  | 2025-05-23 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
365 | Supported Kayne's programming issue. | Genx12 sustenance |  | Completed | 100 | Anubhav, Somesh, Nikhil | N.A. |  |  | 2025-05-23 |  | 2025-05-23 | [5 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
366 | Genx12 Printer stuck Field Issue (Due to paper Guide Incorrect position) | Genx12 sustenance | Field Issue | Completed | 100 |  | Not started | High |  | 2025-06-18 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
367 | Raise Genx12 ECR (Printer Sticker) | Genx12 sustenance | Genx12 ECR(Sticker & Connector) | In progress | 100 | Somesh Sanjana Samiksha Ankar |  | Medium |  | 2025-06-26 | 2025-06-27 | 2025-07-04 | [5 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
368 | ECR was combined with ECR-469-2-25 | Genx12 sustenance | Genx12 ECR(Sticker & Connector) | Completed | 100 | Somesh | Not started | Medium |  | 2025-06-25 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
369 | FMI/FRU will be generated by IB | Genx12 sustenance | Genx12 ECR(Sticker & Connector) | Not started | 100 | Sumesh | Not started | Medium |  | 2025-06-26 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
370 | In BOM (alternate for connector) will be added along with the SAP code (if required). | Genx12 sustenance | Genx12 ECR(Sticker & Connector) | In progress | 100 | Somesh | Not started | Medium |  | 2025-06-26 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
371 | Lijo shall raise ECR for Carton update | Genx12 sustenance | Genx12 ECR(Sticker & Connector) | Not started | 100 | Somesh | Not started | Medium |  | 2025-06-27 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
372 | Get Carton Sample and verify with V&V team. | Genx12 sustenance | Genx12 ECR(Sticker & Connector) | Not started | 100 | Somesh | Not started |  |  | 2025-06-27 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
373 | Create V&V request form. | Genx12 sustenance | Genx12 ECR(Sticker & Connector) | Not started | 100 |  | Not started | Medium |  | 2025-06-27 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
374 | Genx12 Print Instruction Label(Paper Guide) ECR shall be raised. | Genx3 Sustenance | Genx12 ECR(Sticker & Connector) | In progress | 100 |  | Not started | Medium |  | 2025-07-08 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
375 | Protective Film ECR shall be raised or included in other ECR. | Genx12 sustenance | Genx12 ECR(Sticker & Connector) | Not started | 100 |  | Not started | Medium |  | 2025-07-08 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
376 | 2_12Add an option to include or exclude RV5/SV1 and RV5+SV1 information in the auto-generated ECG report | Genx12 sustenance | Genx12 Product Team Input | Completed | 100 | Sanjana | Completed | Medium |  | 2025-04-09 01:00 | [5 empty] | Recommendation from sales Team(GRI Recommendation) Ref Email Subject: Discussion on Recommended ECG Format from Mr. Peter [Glasgow] Dated 1/04/2025 |  | Item | sites/ECGTeam761/Lists/Genx sustenance
377 | (Product team May 2025) Add P/QRS/T durations in minimal report. --Product team In the short and minimal report, P/QRS/T durations are missing and should be included. | Genx12 sustenance | Genx12 Product Team Input | Completed | 100 | @mahesh | Completed | Medium |  | 2025-05-20 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
378 | (Product team May 2025) The parameters and their values at the top of the report needs to be aligned properly. ---Product team (May 2025) | Genx12 sustenance | Genx12 Product Team Input | Completed | 100 | @Mahesh | Completed | Medium |  | 2025-05-20 | 2025-05-27 | 2025-05-30 | [5 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
379 | (Product team May 2025) Option to individually Enable and disable demographic settings shall be provided. And if no field is present, then don't print that option. | Genx12 sustenance | Genx12 Product Team Input | Completed | 100 | Sanjana | Completed |  |  | 2025-05-23 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
380 | (Product team May 2025 -> Outcome) Provide option to include age in months for neonate patients. | Genx12 sustenance | Genx12 Product Team Input | Completed | 100 | Somesh | Completed |  |  | 2025-05-23 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
381 | (Product team May 2025) Message shall be updated "Lead Off Detected" "Connect leads properly and try again for Print OUT | Genx12 sustenance | Genx12 Product Team Input | Completed | 100 | Mahesh D | Completed | Medium |  | 2025-06-06 | 2025-06-09 |  |  |  | 1.check all mode and trigger sample 2.Rt sample and Periodic sample are in same working scenario 3. check presample - work 4.print option for LeedOff Condition 5. store the Record and check all condition ..and increment the patient ID 6after storing record - check the pagesave -periodic - Preview | Lead off Detect - print out |  | Item | sites/ECGTeam761/Lists/Genx sustenance
382 | (Product team May 2025) Remove ST-Duration only from 1 page report (9 line). | Genx12 sustenance | Genx12 Product Team Input | Completed | 100 | Mahesh DE | Completed | Low |  | 2025-05-23 | 2025-06-27 | [6 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
383 | (Product team May 2025) Update the 10/5 Gain Mark as per predicate device. | Genx12 sustenance | Genx12 Product Team Input | Completed | 100 |  | Completed |  |  | 2025-05-29 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
384 | Unit Test 1 By Samiksha for 1.01 | Genx12 sustenance | Genx12 Product Team Input | Completed | 100 | Samiksha | Not started | High | 2025-07-24 17:00 | 2025-07-28 | 2025-07-01 | 2025-07-24 | Samiksha |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
385 | Genx12 Unit Test By Somesh | Genx12 sustenance | Genx12 Product Team Input | Completed | 100 | Somesh |  | High |  | 2025-07-28 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
386 | Long Xc and TDK lambda SMPS Genx12 EMI EMC testing and ECR(ECR-469-02-2025) closure. | Genx12 sustenance | Genx12-ECR(SMPS) | In progress | 100 | someshwaran.s@bpl.in | N.A. | Medium |  | 2025-05-14 | 2025-05-15 | 2025-05-23 |  |  | Will be attached as an attachment GENX12 ECR-ECN TASKS.xlsx |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
387 | SMPS LongXc FRU (Prepare) | Genx12 sustenance | Genx12-ECR(SMPS) | Completed | 100 | Somesh | N.A. | High | 2025-06-05 16:00 | 2025-06-02 |  |  | Somesh |  |  |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
388 | TDK lambda FRU (Prepare) | Genx12 sustenance | Genx12-ECR(SMPS) | Completed | 100 | Somesh | N.A. | High |  | 2025-06-02 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
389 | TDK lambda BOM UPDATE | Genx12 sustenance | Genx12-ECR(SMPS) | Completed | 100 | Somesh | N.A. | High |  | 2025-06-03 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
390 | TDK lambda SAP CODE (CREATE) | Genx12 sustenance | Genx12-ECR(SMPS) | Completed | 100 | Somesh | N.A. | High |  | 2025-06-03 00:11 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
391 | TDK lambda Service Manual (Update) | Genx12 sustenance | Genx12-ECR(SMPS) | Completed | 100 | Somesh | N.A. | High |  | 2025-06-03 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
392 | TDK lambda (Cable drawing update) | Genx12 sustenance | Genx12-ECR(SMPS) | Completed | 100 | Somesh | N.A. | High |  | 2025-06-03 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
393 | TDK lambda spare part update | Genx12 sustenance | Genx12-ECR(SMPS) | Completed | 100 | Somesh | N.A. | High |  | 2025-06-03 00:13 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
394 | TDK lambda Board Alignment procedure update | Genx12 sustenance | Genx12-ECR(SMPS) | Completed | 100 | Somesh | N.A. | High |  | 2025-06-03 00:14 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
395 | ECR(ECR-469-02-2025) V & V testing report. | Genx12 sustenance | Genx12-ECR(SMPS) | In progress | 100 | Nikhil | N.A. | High |  | 2025-06-03 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
396 | SMPS ECN Impact Analysis (Create) | Genx12 sustenance | Genx12-ECR(SMPS) | Completed | 100 | Somesh | N.A. | High |  | 2025-06-03 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
397 | Paper Roll Carton (SAP code update) | Genx12 sustenance | Genx12-ECR(SMPS) | Completed | 100 |  | N.A. | High |  | 2025-06-03 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
398 | paper roll carton BOM update | Genx12 sustenance | Genx12-ECR(SMPS) | Completed | 100 |  | N.A. | High |  | 2025-06-03 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
399 | BIS update | Genx12 sustenance | Genx12-ECR(SMPS) | HOLD(Dependency) | 100 |  | N.A. | High |  | 2025-06-03 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
400 | DTL and DOF update | Genx12 sustenance | Genx12-ECR(SMPS) | Completed | 100 | Samiksha and Somesh | Not started | High |  | 2025-06-10 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
401 | Genx3 Filter Implementation (Same as Genx12) | Genx3 Sustenance | Genx3_ECR(SMPS) | Not started | 100 |  | N.A. | Low |  | 2025-05-26 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
402 | Reopen ECR | Genx3 Sustenance | Genx3_ECR(SMPS) | Completed | 100 | Somesh | Not started | Medium |  | 2025-06-09 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
403 | CCL (SMPS + Neonatal Cable) | Genx3 Sustenance | Genx3_ECR(SMPS) | Completed | 100 | Somesh | Not started | Medium |  | 2025-06-09 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
404 | BOM Update SAP code, Neonatal Cable, SMPS Assembly | Genx3 Sustenance | Genx3_ECR(SMPS) | Completed | 100 | Sanjana | Not started | Medium |  | 2025-06-09 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
405 | Create Spare part list | Genx3 Sustenance | Genx3_ECR(SMPS) | Completed | 100 | Sanjana Sampige | Not started | Medium |  | 2025-06-09 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
406 | Create Cable drawing for Genx3i and Genx3.(SMPS EMERSON (NPS44-M)& LongXc (ACMS 43-150)) | Genx3 Sustenance | Genx3_ECR(SMPS) | Completed | 100 | Sanjana | Not started | Medium |  | 2025-06-09 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
407 | Update Service Manual (Neonatal Cable) | Genx3 Sustenance | Genx3_ECR(SMPS) | In progress | 100 | Sanjana | Not started | Medium |  | 2025-06-09 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
408 | Update User Manual (Neonatal Cable) | Genx3 Sustenance | Genx3_ECR(SMPS) | Completed | 100 | Samiksha | Not started | Medium |  | 2025-06-09 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
409 | Create ECN Empact Analysis | Genx3 Sustenance | Genx3_ECR(SMPS) | Completed | 100 | Sanjana | Not started | Medium |  | 2025-06-09 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
410 | Create SMPS FRU Genx3 and Genx3i | Genx3 Sustenance | Genx3_ECR(SMPS) | Completed | 100 | Sumesh and Somesh | Not started | Medium |  | 2025-06-09 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
411 | Internal Test Report (SMPS) | Genx3 Sustenance | Genx3_ECR(SMPS) | Completed | 100 | Nikhil | Not started | Medium |  | 2025-06-09 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
412 | Create Internal test report Neonatal Cable | Genx3 Sustenance | Genx3_ECR(SMPS) | Completed | 100 | Beryl | Not started | Medium |  | 2025-06-10 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
413 | Create V&V Request Form | Genx3 Sustenance | Genx3_ECR(SMPS) | Completed | 100 | Samiksha | Not started | Medium |  | 2025-06-10 |  | 2025-06-26 | [5 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
414 | Machine Dispatch from Astute | Genx3 Sustenance | Genx3_ECR(SMPS) | Completed | 100 | Aman Singh Negi | Not started | Medium |  | 2025-06-10 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
415 | Update DOF and DTL | Genx3 Sustenance | Genx3_ECR(SMPS) | Completed | 100 |  | Not started | Medium |  | 2025-06-10 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
416 | Neonate Cable Packing Sticker or carton (FEASIBILITY, Finalization and implementation) | Genx3 Sustenance | Genx3_ECR(SMPS) | In progress | 100 |  | Not started | High |  | 2025-07-02 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
417 | Add Genx3 Alternate component to ECR (SMPS) | Genx3 Sustenance | Genx3_ECR(SMPS) | In progress | 100 |  | Not started | Medium |  | 2025-07-08 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
418 | Genx3 Assembly Issue Tracker needs to be addressed | Genx3 Sustenance | Genx3-PILOT | In progress | 100 | Someshwaran S | false | Medium |  | 2025-03-18 01:00 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
419 | Give two units to V&V for Lonxc and emerson SMPS . | Genx3 Sustenance | Genx3-PILOT | Completed | 100 | Owner: Somesh, chandan kumar | false | Medium |  | 2025-03-03 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
420 | Genx3 CE production hardware Support (PCB to FG) | Genx3 Sustenance | Genx3-PILOT | In progress | 100 | Nikhil, Somesh, Nitin and Anubhav | N.A. | High |  | 2025-05-09 | 2025-04-28 | 2025-05-14 | [5 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
421 | Genx3 CE BOM finalization | Genx3 Sustenance | Genx3-PILOT | In progress | 100 | Somesh and Nitin | N.A. | High | 2025-05-15 | 2025-04-28 | 2025-04-28 | 2025-05-15 | [5 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
422 | Shield ECG Board-(Sample Expected by 6 June) | Genx3 Sustenance | Genx3-PILOT | In progress | 100 |  | N.A. | High |  | 2025-06-02 |  | 2025-06-06 | [5 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
423 | Main Board Myler Sheet (Sample will be received by 4 June). | Genx3 Sustenance | Genx3-PILOT | In progress | 100 |  | N.A. | High |  | 2025-06-02 |  | 2025-06-04 | [5 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
424 | SMPS-Emerson and Long Xc (Check Cable for both SMPS). | Genx3 Sustenance | Genx3-PILOT | In progress | 100 | Somesh | N.A. | High |  | 2025-06-02 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
425 | Write email deviation note or instruction for Cable (Length + ferrite core). | Genx3 Sustenance | Genx3-PILOT | Completed | 100 | Anubhav | N.A. | High |  | 2025-06-02 |  | 2025-06-03 | [5 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
426 | Screen Printing (02/06/2025) | Genx3 Sustenance | Genx3-PILOT | Completed | 100 |  | N.A. | High |  | 2025-06-02 |  | 2025-06-02 | [5 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
427 | Machine Assembly for Pilot. (After all parts received) | Genx3 Sustenance | Genx3-PILOT | In progress | 100 |  | N.A. | High |  | 2025-06-02 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
428 | V&V AND ECN Closure. Raise ECR-by SCM . Work on ECN . User Manual will be updated. Service Manual will be updated. Spare part list will be updated. | Genx3 Sustenance | Genx3-PILOT | In progress | 100 |  | N.A. | High |  | 2025-06-02 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
429 | Review Genx3 Draft Report from Astute Lab | Genx3 Sustenance | Genx3-PILOT | In progress | 100 | Anubhav | N.A. | High |  | 2025-06-03 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
430 | Genx3 Top and Base Closure issue | Genx3 Sustenance | Genx3-PILOT | In progress | 100 | Somesh, Anubhav | Not started | High |  | 2025-06-17 | 2025-06-13 | 2025-06-19 | [5 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
431 | 5_12. 3x4+3R Format Issue: In the 3x4+3R format, the last rhythm lead is shifted towards the footer of the page. | Genx12 sustenance | Miscellaneous | Not started | 100 | Mahesh DE | false | Medium |  | 2025-04-09 01:00 | [5 empty] | Option3: Reduce the content in header permanently for one line summary and then print. Ref: Re: Observations from ECG Camp – GenX12i+ Performance and Improvement Suggestions 7/03/2025 |  | Item | sites/ECGTeam761/Lists/Genx sustenance
432 | 4_18 ECG trace shall start from the gain mark | Genx12 sustenance | Miscellaneous | Not started | 100 | Sanjana | false | Low |  | 2025-04-09 01:00 | [5 empty] | Recommendation from sales Team(GRI Recommendation) Ref Email Subject: Discussion on Recommended ECG Format from Mr. Peter [Glasgow] Dated 1/04/2025 |  | Item | sites/ECGTeam761/Lists/Genx sustenance
433 | Include Cyber security Hazards in RMR File. MDR asks for this. | Genx12 sustenance | Miscellaneous | Not started | 100 | Somesh | false | Medium |  | 2025-04-07 01:00 | [5 empty] | Copy Hazards from Genx12 Cyber security report to RMR. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
434 | Printer Noise reduction create Suggestion | New Platform Tasks | Miscellaneous | Not started | 100 | Anubhav Mahesh Somesh Sanjana | false | Medium |  | 2025-04-02 01:00 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
435 | How Lead off works in 9108,GE and mindray machine. | New Platform Tasks | Miscellaneous | Not started | 100 | Owner: Sanjana , Somesh ,Anubhav ,Mahesh | false | Medium |  | 2025-04-02 01:00 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
436 | With One Line Summary reduce the header and shift ECG trace UP. | Genx12 sustenance | Miscellaneous | Not started | 100 | Mahesh, Nitin | false | Low |  | 2025-03-24 01:00 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
437 | Scope for SECURE FTP(SFTP) in our Genx3 and Genx12 device. | Genx12 sustenance | Miscellaneous | Completed | 100 | Anubhav | false | Medium |  | 2025-03-12 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
438 | RT View REQ : Name ID addition ,time stamp and initials shall be part of folder and file name | Genx12 sustenance | Miscellaneous | Not started | 100 | Anubhav Chaudhary | false | Medium |  | 2025-02-18 | [5 empty] | We have this info in our report name TAG So this requirement is complete. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
439 | RT View REQ: Asses where to add Patient Demographics for PDF for FTP upload. | Genx12 sustenance | Miscellaneous | Not started | 100 | Anubhav Chaudhary | false | Medium |  | 2025-02-18 |  |  |  | 1/1 | Assesment Done | Option 1: As per my suggestion it shall be read from PDF file by the PC based software by creating INward and outward folders. Option 2: is sending other CSV file before PDF transmission with these demo graphic details. Draw back of sending CSV file: If file is missed then there wont want be no info(Or info will be missed). It will add extra time for transmission. Option 3: Just send HL7 file on FTP to server.(HL7 file has all data). Once file reached it will go to INWARD folder and then converted desired PDF along with expected data automatically. |  | Item | sites/ECGTeam761/Lists/Genx sustenance
440 | Genx3 CE Validation report and verification separate doc is required. CREATE ONE | Genx3 Sustenance | Miscellaneous | Not started | 100 |  | false | Medium |  | 2025-02-17 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
441 | Predicate Device Feature observation | New Platform Tasks | Miscellaneous | Not started | 100 | Owner:Sanjana, Somesh, Anubhav, Mahesh | false | Medium |  | 2025-02-04 | 2025-02-04 |  |  |  |  | GE MAC 5 Basic operation duration |  | Item | sites/ECGTeam761/Lists/Genx sustenance
442 | Sometimes Lead off indication is observed in FREEZE MODE | Genx12 sustenance | Miscellaneous | Not started | 100 | Mahesh DE | false | Medium |  | 2025-01-31 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
443 | Genx12 RT View. all record transfer issue | Genx12 sustenance | Miscellaneous | In progress | 100 | Anubhav Chaudhary;Someshwaran S;sanjana sampige | false | Medium |  | 2025-01-27 |  | 2025-03-20 01:00 |  | 5/8 | Concept tested on pseudo code.;Integration for Genx3 (Implementation);Genx3 testing for the change.;Higher baud rate feasibility.;Integration for Genx12 (Implementation);Issue Assessment Done;Planned new logic;Genx12 testing for the change. | Tested with Real Term utility (Char delay is set as 10 msec) Ack packet delay is kept as 2000 msec(2 Sec) Ajay(From S/W) modified the software for 10 msec but still records not transferred. As per Somesh Observation with modified software data without interpretation is not getting transferred. But old software it is working fine. 03-2-25 observations: Char delay-10m/s and send in send (repeat delay)-1000m/s = All the records in the machine transferred. Char delay-5/s and send in send (repeat delay)-1000m/s = All the records in the machine transferred. Char delay-2.5m/s and send in send (repeat delay)-1000m/s = All the records in the machine transferred. Char delay-2m/s and send in send (repeat delay)-1000&2000m/s = All the records in the machine transferred for one time ,second time it only transferred 36 records Char delay-1.5/s and send in send (repeat delay)-1000m/s = All the records in the machine transferred. Each record taking 18 secs (approx.) to transfer from machine |  | Item | sites/ECGTeam761/Lists/Genx sustenance
444 | Update Genx3 Brochure with latest interpretation images. | Genx3 Sustenance | Miscellaneous | Not started | 100 | samiksha shetty | false | Medium |  | 2025-01-21 |  |  |  | 0/1 | Collect images from Sindhu and initiate mail with marketing team. |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
445 | User manual study | New Platform Tasks | Miscellaneous | Not started | 100 |  | false | Medium |  | 2025-01-20 |  |  |  | 0/1 | GE Machine User manual Study |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
446 | platform Study | New Platform Tasks | Miscellaneous | Not started | 100 |  | false | Medium |  | 2025-01-20 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
447 | usability study | New Platform Tasks | Miscellaneous | Not started | 100 |  | false | Medium |  | 2025-01-20 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
448 | User Interface | New Platform Tasks | Miscellaneous | Not started | 100 |  | false | Medium |  | 2025-01-20 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
449 | Study Features | New Platform Tasks | Miscellaneous | Not started | 100 |  | false | Medium |  | 2025-01-20 |  |  |  | 0/1 | Filter Range Study |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
450 | Genx3 Brochure Print image update | Genx12 Design Transfer | Miscellaneous | Not started | 100 | Someshwaran S | false | Medium |  | 2025-01-11 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
451 | PDF Comparison of older Dot Net version with Latest One. | Genx12 sustenance | Miscellaneous | Not started | 100 |  | false | Medium |  | 2024-11-05 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
452 | Map Lead Off indication with LCD trace ,Print Trace and PC trace | Genx12 sustenance | Miscellaneous | Not started | 100 |  | false | Low |  | 2024-09-11 01:00 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
453 | 2.1 Printer noise detected. | Genx12 sustenance | Miscellaneous | Not started | 100 |  | false | Medium |  | 2024-09-07 01:00 |  |  |  | 0/2 | Simulation required to identify noise.;• Need Q/A observation for the deviation. |  |  | Item | sites/ECGTeam761/Lists/Genx sustenance
454 | auto-Trigger Mode-Pacemaker Disable Message added | Genx12 sustenance | Miscellaneous | In progress | 100 | Mahesh DE | N.A. | Medium |  | 2025-05-14 | 2025-05-15 |  | Mahesh DE | check all settings | At trigger mode - there was showing message as ENABLE THE INTERPRETATION, -Interpretation enable also Because Pacemaker -ON, so Added Message as Disable PACEMAKER at patient Info | MESSAGE IS AS "Please Enable Interpretation/Measurement by press Menu Key and Select Interpretation. And Disable the pacemaker by pressing Patient Info and F! |  | Item | sites/ECGTeam761/Lists/Genx sustenance
455 | Check the all freeze format and print format and Display format | Genx12 sustenance | Miscellaneous | In progress | 100 | mahesh DE | N.A. | Medium |  | 2025-05-20 | 2025-05-20 | 2025-05-22 |  |  | auto mode-Check print format- display format and freeze format manual mode - Check print format- display format and freeze format page save-Check print format- display format and freeze format | Main Observation : -check pages and print format - check any mismatches in Freeze condition -check the spike on display /print - check the print format mismatch EX- if print format is 6x2 but printing as 3x4 - check the remain home screen display format and print format |  | Item | sites/ECGTeam761/Lists/Genx sustenance
456 | Fijitsu 50mm speed Isuue observed(After printing first page Motor not running humming sound is comming) | Genx12 sustenance | Miscellaneous | In progress | 100 | Mahesh | Not started |  |  | 2025-07-18 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
457 | Plan and Add login Credentials to access Genx12 and Genx3 ECG rec and connectivity section | Genx12 sustenance | Miscellaneous | Not started | 100 |  | Not started | Low |  | 2025-06-29 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
458 | Check feasibility and implement DIACOM on Dot Net Board. | Genx12 sustenance | Miscellaneous | Not started | 100 | Sanjana | Not started | Low |  | 2025-06-30 | [7 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance
459 | Lead OFF and Skin Impedance Study | Feasibility study | Study Analysis Feasibility | In progress | 25 | Sanjana and Samiksha | N.A. | Medium |  | 2025-05-09 | 2025-05-01 | 2025-06-06 | [5 empty] | Item | sites/ECGTeam761/Lists/Genx sustenance